## Slide 1: Paul Scherrer Institute, Switzerland
Front-end Electronics, Waveform Digitizing and Signal Processing
Stefan Ritt

## Slide 2: “Manual” DAQ
1000 tracks per 25 ns
A4 paper @ 5 g
Truck load @ 40 t
How much paper
per second?
200’000 t
5000 Trucks !!!

## Slide 3: Electronic Signal Acquisition
Particle
Waveform Digitizing and Signal Processing
Electronic Signal
ADC & TDC technologies
Signal shaping
Ultra-fast digitizing (>1 GSPS)
Digital pulse processing
Applications

## Slide 4: Signals in particle physics
Photomultiplier (PMT)
Scintillator
Particle
10 – 100 ns
HV
HV
1 – 10 ms
Scintillators
(Plastic, Crystals, Noble Liquids, …)
Wire chambers Straw tubes
Silicon
Germanium

## Slide 5: Measure precise timing: ToF-PET
Positron Emission Tomography
Time-of-Flight PET
Dt
d
d ~ c/2 * Dt
e.g. d=1 cm → Dt = 67 ps

## Slide 6: 1-bit & 2-bit ADC
Threshold
VREF
VIN
B
1-Bit Flash
Analog to Digital
Converter
R
R

## Slide 7: n-bit ADC
Flash ADC very fast for small number of bits
Requires 2n comparators

## Slide 8: Successive approximation ADC

## Slide 9: Pipeline ADC
Combine 4-Bit flash ADC with successive approximation logic
Only requires 4-Bit flash ADC
Can convert one sample in each clock cycle
Has a latency depending on the number of pipeline stages
Most common technology for fast ADCs (> 10 MHz)

## Slide 10: Digital-to-analog converter (DAC)
R-2R Ladder: 0/VR at S0...Sn-1 give binary weighted output Vo = S * VR/32
Sn-1-S0 | VO [V]
00000 | 0
00001 | 0.16
00010 | 0.31
00011 | 0.47
00100 | 0.63
00101 | 0.78
... | ...
11100 | 4.38
11101 | 4.53
11110 | 4.69
11111 | 4.84
5 V

## Slide 11: Raspberry Pi ADC – DAC Board
ADC
DAC

## Slide 12: Time-to-Digital Converter
Dt
Coarse Counter
Stop
Clock
TDC
Coarse
Output
TDC
Fine
Output
1
1
0
0
0

## Slide 13: TDC in FPGA
Real Time 2016 Conference, Padova

## Slide 14: Signal discrimination
Threshold
Single Threshold
“Time-Walk”
Multiple Thresholds
T1
T2
T3
T1
T2
T3
Inverter & Attenuator
S
Delay
Adder
0
Constant Fraction (CFD)

## Slide 15: Influence of noise
Voltage noise causes
timing jitter !
Fourier Spectrum
Signal
Noise
Low pass filter
Low pass filter (shaper) reduces noise while maintaining most of the signal

## Slide 16: Noise limited time accuracy
U [mV] | DU [mV] | tr | Dt
100 | 1 | 1 ns | 10 ps
10 | 1 | 3 ns | 300 ps
All values in this talk are s (RMS) !
FHWM = 2.35 x s
Most today’s TDCs have ~20 ps LSB
How can we do better ?

## Slide 17: Noise limited time accuracy

## Slide 18: Example: CFG in FPGA
+
Adder
Look-up Table
(LUT)
8-bit address
8-bit data
* (-0.3)
S
Delay
Adder
0
Latch
Clock
>0
≤ 0
AND
Delay
FPGA

## Slide 19: Nyquist-Shannon Sampling Theorem
fsignal < fsampling /2
fsignal > fsampling /2

## Slide 20: Limits of waveform digitizing
Aliasing Occurs if fsignal > 0.5 * fsampling
Features of the signal can be lost (“pile-up”)
Measurement of time becomes hard
ADC resolution limits energy measurement
Need very fast high resolution ADC

## Slide 21: Charge ADCs (QADC)
Threshold
Threshold
TDC
(Clock)
+
-
ADC
~MHz

## Slide 22: Peak sensing ADC
Threshold
Threshold
Delay
+
-
ADC
~MHz
Sample

## Slide 23: Double buffering
ADC conversion time
+
-
ADC
~MHz
+
-
Store an event in an analog cell
Start converting cell
If second event arrived before conversion is ready – store event in second cell
Decrease dead time significantly

## Slide 24: Analog readout chain
Effect of slow shaping time:
“pile-up”
Optimal parameters can greatly improve the signal-to-noise ratio

## Slide 25: Baseline noise
Baseline with ~50 Hz noise
Charge integration error due to signal “sitting” on fluctuating baseline (e.g. 50 Hz ground loop or artifact of shaper)
Can be fixed by sampling baseline prior to signal (requires signal delay)
Sample signal after peak for pile-up recognition

## Slide 26: Conventional DAQ chain
Digital world
Bits and Bytes
Detector world
Energy, Time
Analog world
Charges, Voltages
Particle
Detector
Signal
Conditioning
Cable
Analog-to-Digital
Converter
Processing
Scintillator
Solid state detector
Gas detector
Amplifier
Shaper
Discriminator
Digital filter
Charge integrator
Waveform fitting
High-level triggering
Data storage

## Slide 27: Let’s go a bit faster…

## Slide 28: Ultimate sampling
Direct fast sampling without shaping
  No shaping artifacts
  Less electronics
  All information if captured if fsampling > 2*fsignal and LSB < Vnoise
  Any shaping circuitry can only remove information
Fast ADC

## Slide 29: What are the fastest detectors?
Micro-Channel-Plates (MCP)
      Photomultipliers with thousands of tiny channels (3-10 mm)
      Typical gain of 10,000 per plate
      Very fast rise time down to 70 ps
70 ps rise time  4-5 GHz BW  10 GSPS
SiPMs (Silicon PMTs) are also getting < 100 ps
J. Milnes, J. Howoth, Photek
http://sensl.com

## Slide 30: Can it be done with FADCs?
8 bits – 3 GS/s – 1.9 W  24 Gbits/s
10 bits – 3 GS/s – 3.6 W  30 Gbits/s
12 bits – 3.6 GS/s – 3.9 W  43.2 Gbits/s
14 bits – 0.4 GS/s – 2.5 W  5.6 Gbits/s
1.8 GHz!
24x1.8 Gbits/s
Requires high-end FPGA
Complex board design
High FPGA power
PX1500-4: 2 Channel 3 GS/s 8 bits
ADC12D1X00RB: 1 Channel 1.8 GS/s 12 bits
1-10 k$ / channel
What about 1000+ Channels?
V1761: 2 Channels, 4 GS/s, 10 bits

## Slide 31: Switched Capacitor Array (Analog Memory)
Shift Register
Clock
IN
Out
“Time stretcher” GHz  MHz
Waveform
stored
Inverter “Domino” ring chain
0.2-2 ns
FADC
33 MHz
10-100 mW

## Slide 32: Triggered Operation
sampling
digitization
lost events
sampling
digitization
Sampling Windows * TSR
Chips usually cannot sample during readout ⇒ “Dead Time”
Technique only works for “events” and “triggers”
Dead time =
Conversion Time * Samples
(e.g. 30 ns ∙ 100 = 3 ms)

## Slide 33: How to measure best timing?
Simulation of MCP with realistic noise and different discriminators
J.-F. Genat et al., arXiv:0810.5590 (2008)
D. Breton et al., NIM A629, 123 (2011)
Beam measurement at SLAC & Fermilab

## Slide 34: How is timing resolution affected?
voltage noise Du
timing uncertainty Dt
signal height U
rise time tr
number of samples on slope
Simplified estimation!

## Slide 35: How is timing resolution affected?
U | Du | fs | f3db | Dt
100 mV | 1 mV | 2 GSPS | 300 MHz | ∼10 ps
1 V | 1 mV | 2 GSPS | 300 MHz | 1 ps
1 V | 1 mV | 10 GSPS | 3 GHz | 0.1 ps
today:
optimized SNR:
next generation:
Assumes ideal sampling
“Novel Calibration Method for Switched Capacitor Arrays Enables Time Measurements width Sub-Picosecond Resolution”, D.A. Stricker-Shaver, S. Ritt, B.J. Pichler, IEEE TNS 61 (2014), 3607

## Slide 36: Readout of Straw Tubes
HV
Readout of straw tubes or drift chambers usually with “charge sharing”: 1-2 cm resolution
Readout with fast timing: 10 ps / √10 = 3 ps → 0.5 mm
Currently ongoing research project at PSI
d ~ c/2 * Dt

## Slide 37: Switched Capacitor Arrays for Particle Physics
STRAW3
TARGET
LABRADOR3
AFTER
NECTAR0
SAM
E. Delagnes
D. Breton
CEA Saclay
DRS1
DRS2
DRS3
DRS4
G. Varner, Univ. of Hawaii
0.25 mm TSMC
Many chips for different projects (Belle, Anita, IceCube …)
0.35 mm AMS
T2K TPC, Antares, Hess2, CTA
H. Frisch et al., Univ. Chicago
PSEC1 - PSEC4
0.13 mm IBM
Large Area Picosecond Photo-Detectors Project (LAPPD)
2002
2004
2007
2008
0.25 mm UMC
Universal chip for many applications
MEG experiment, MAGIC, Veritas, TOF-PET
SR
R. Dinapoli
PSI, Switzerland
drs.web.psi.ch
www.phys.hawaii.edu/~idlab/
matacq.free.fr
psec.uchicago.edu
Poster 15, 106

## Slide 38: MEG On-line waveform display
template
fit
S848
PMTs
“virtual oscilloscope”
Liq. Xe
PMT
1.5m
g
m+e+g
At 10-13 level
3000 Channels
Digitized with DRS4 chips at
1.6 GSPS
m
Drawback: 400 TB data/year
750 Oscilloscopes!

## Slide 39: Pulse shape discrimination
g
a
m
g
a
m
Events found and correctly processed 2 years (!) after the were acquired

## Slide 40: MAGIC Telescope
http://ihp-lx.ethz.ch/Stamet/magic/magicIntro.html
https://wwwmagic.mpp.mpg.de/
La Palma, Canary Islands, Spain, 2200 m above sea level

## Slide 41: Template Fit
Determine “standard” PMT pulse by averaging over many events  “Template”
  Find hit in waveform
  Shift (“TDC”) and scale (“ADC”) template to hit
  Minimize c2
  Compare fit with waveform
  Repeat if above threshold
Store ADC & TDC values
pb Experiment
500 MHz sampling
www.southerninnovation.com
14 bit
60 MHz

## Slide 42: High speed USB oscilloscope
4 channels
5 GSPS
1 GHz BW
8 bit (6-7)
15 k€
4 channels
5 GSPS
1 GHz BW
11.5 bits
1170 €
USB Power

## Slide 43: New chips
Self-trigger writing of 128 short 32-bin segments (4096 bins total)
Storage of 128 events
  Accommodate long trigger latencies
  Quasi dead time-free up to a few MHz,
  Possibility to skip segments → second level trigger
SAMPIC Waveform TDC
Record short (64 bins) waveforms
Digitize on-chip
Data-driven read out
DRS5 (PSI, planned)
CEA/Saclay

## Slide 44: Summary
Digitization is a key element of all particle physics experiments
General trend to faster digitization and waveform analysis in digital domain (embedded CPUs, FPGAs)
This talk can only give you a glimpse
Further information
  H. Spieler, “Semiconductor Detector Systems”, Oxford Univ. Press, 2005
  G. Knoll, “Radiation Detection and Measurement”, Wiley, 2010
  Conferences (with short courses): IEEE NSS-MIC (Sydney, Australia, Nov. 2018) IEEE Realtime (Ho Chi Minh City, Vietnam, June 2020)
  Become IEEE NPSS member