## Slide 1: Semiconductor hybrid pixel detectors for detection of neutrons
Stanislav Pospíšil
Institute of Experimental and Applied Physics
Czech Technical University in Prague
https://indico.cern.ch/event/954199/

## Slide 2: Outline
Efficient position and energy sensitive detection of neutrons is fundamental for advance in neutron imaging, neutron dosimetry and within many other subfields of research, where neutrons are used.
The general principles of neutron detection based on a brief overview of specific interactions of neutrons of different energies will be reminded in the presentation.
The presentation will focus on the concept of semiconductor pixel detectors with converters of neutrons to strongly ionizing charged particles designed for detection of neutrons in radiation hard environment.
Particular attention will be given to challenges associated with growing interests in neutron imaging with submicron resolution and in noiseless detection of single neutrons in mixed n-gamma radiation fields.
Perspectives of the proposed concept will be backed up by the latest experimental results obtained with Timepix pixel detectors using the methodology of "particle track pattern recognition".

## Slide 3: The world of neutrons is very colorful
It deals with neutrons of kinetic energies Tn in the range of 15 orders, 10-7 eV <Tn <108 eV, when we talk about ultra-cold, cold, thermal, resonance, intermediate, fast and relativistic neutrons.
In all reactions used to release neutrons from nuclei, born neutrons are always fast. Slowing down processes based on neutron scattering with nuclei are responsible for change of their energies in all environment.
The nature of neutron interactions with nuclei depends significantly on Tn:
In the case of cold, thermal and resonance neutrons, the neutron wavelengths have to be taken in account and radiative capture of neutrons (n,γ) is dominating process.
In the case of intermediate and fast neutrons, strong interactions with atomic nuclei dominate as two-particle interactions, such as elastic scattering and inelastic scattering, and classical (n, alpha) or (n, p) nuclear reactions.
In the case of neutrons with energies higher than 20 MeV, more complex reactions and fragmentation processes are becoming crucial.

## Slide 4: About neutron converters, detectors and radiation background
All of the above mentioned processes can be used to detect neutrons. However, the choice of the converter depends on what neutrons are to be detected and for what purpose the detectors are to be used. This applies to all types of detectors, be the gas, scintillation (liquid, solid, hydrogen rich) or semiconductor detectors.
Neutrons from the source are always accompanied by high-energy gamma radiation from inelastic scattering, nuclear reactions and above all from radiation capture of neutrons (with energies up to 11 MeV) on the nuclei of elements from the surrounding environment.
Indeed, one can find everywhere a number of nice publications describing neutron detectors with high neutron detection efficiency, which however suffer from high efficiency also to gamma rays. Then, to get a net neutron-induced signal only, the undesirable gamma signal has to be discriminated, what is not an easy task at all. However, this problem can be successfully overcome by using a thin sensor with the neutron converter to energetic ions.

## Slide 5: Why semiconductor neutron detectors can compete with gas and scintillating detectors?
Silicon detectors had become well developed for high resolution alpha and energetic ions spectroscopy in nuclear physics and for many applications profiting from their excellent resolution, window less detection, low bias, low power, small dimensions opening possibility to make portable devices.
The adaption of the detectors with small sensitive volume for detection and spectroscopy of neutrons based on their conversion to highly ionizing ions of relatively high energies permits the reliable discrimination of neutrons from gamma rays and electrons. Such devices were optimized for monitoring of neutrons in mixed n-gamma radiation fields with applications in nuclear reactor experiments and different moisture meters.
In the case of pixel detectors with neutron converters to short-range energetic particles (like 6LiF, 10B, hydrogen rich material as CH2, 235,238U), a thin pixelated sensor allows then discrimination of unwanted signal from gamma photons according to energy, but also "noiseless" recognition of a neutron according to its track observed in the detector.

## Slide 6: Response of the silicon diode with CH2 converter to 4.2 MeV neutrons (recoil proton spectroscopy)
Silicon diode with CH2 converter
Silicon diode without CH2 (background)
Pure spectrum of recoiled protons
Pulse height spectrum of Silicon diode when illuminated by fast neutrons (14.8 MeV) at different bias (0, 5, 20 V). Peaks from interaction of fast neutrons with 28Si are clearly seen

## Slide 7: Responses of semiconductor diode with 6LiF converter to thermal and fast neutrons
Silicon diode illuminated
by fast neutrons
Pulse height distributions of neutron detector on thermal neutrons at different bias (0, 5, 20 V). Detector operates well in the self-biased regime.
Pulse height distributions of neutron detector when illuminated by fast neutrons (14.8 MeV) at different bias (0, 5, 20 V). Peaks from interaction of fast neutrons with 28Si are clearly seen
Silicon diode + 6LiF converter
Illuminated by thermal neutrons

## Slide 8: Medipix/Timepix hybrid pixel detector device
Medipix2/Timepix
Pixels: 256 x 256
Pixel size: 55 x 55 mm2
Area: 1.5 x 1.5 cm2
Medipix2/Timepix
Quad
Pixels: 512 x 512
Pixel size: 55 x 55 mm2
Area: 3 x 3 cm2
www.cern.ch/medipix
Planar pixellated detector (Si, GaAs, CdTe, thickness: 300/700/1000mm)
Bump-bonded to Medipix readout chip containing in each pixel cell:
- amplifier,
- double discriminator
- and counter
Detector chip
Medipix chip
Bump-bonding
a, b, g, ...
● Converter materials to detect
- thermal neutrons: 6Li(n,α)T, Q=4.78MeV
10B(n,α)7Li, Q=2.78MeV
- fast neutrons: recoiled protons from PE-foil
Notes: 8

## Slide 9: Medipix – single quantum counting detector Timepix - spectroscopic pixel detector with ToT and ToA modes of operation
N+
P+
Si
Counter:
+
Amplifier
Compa
rator
000
001
Threshold level above
electronic noise
No false counting.
Digital integration (counting)
No dark current.
Unlimited dynamic range
and exposure time. Counts obey poissonian distribution.
65k spectroscopic chains:
SCA in case of Medipix
MCA in case of Timepix
MCA+TDC with Timepix3
Energy calibration (Calibration of 65k MCA! Question: how to deposit defined energy into a volume 55x55x300 μm3?)
Bias Voltage
Common
back-side
electrode
Pixelated
front-side
electrode
Pixel electronics
Threshold
level
007
003
Particle count
Energy
Timepix TOT mode
Notes: 9

## Slide 10: High resolution X-ray radiography Experimental setup
Geometric magnification
X-ray tube
Object
Enlarged projection
Requirements:
Microfocus X-ray source to enable geometrical magnification
Adjustable object holder (three translations + rotation)
Sample stabilization (temperature, humidity)
Equipment for automatic calibration of pixel responses
Detector holder and detector stabilization (temperature, condensing point)
Notes: In this picture there is shown our Experimental setup. On the left side we can see the Hamammatsu x ray tube and the translation stages with a sample. On the opposite side there are the detector Medipix and carroousel.

## Slide 11: Mouse Kidney Tomography
Missing angles => Iterative algorithm instead of Filtered back projection (3 iterations in OSEM 5)
Soft tissue X-ray imaging
Notes: Next, tomography was followed by reconstruction of every slice. For calculation of missing angles we use interactive algorithm instead (insted) of filtered back projection.
OSEM = Oredered Subset Expectancy Maximization (size of subset = 5)

## Slide 12: Motivation – neutron radiography
X-rays
Neutrons
While X-rays are attenuated more effectively by heavier materials like metals, neutrons allow to image some light materials such as hydrogenous substances with high contrast.
Neutron radiography can serve as complementary technique to X-ray radiography
In the X-ray image, the metal parts of the photo camera are seen clearly, while the neutron radiogram shows details of the plastic parts (courtesy of E. Lehmann, PSI)

## Slide 13: The Neutronography based on 6Li(n,α)T, Q=4.78MeV and10B(n,α)7Li, Q=2.78MeV (energetic charged products define interaction with deeply subpixel resolution)
Specimen attenuating the beam
Shadow on detector
plane
Neutronogram
Parallel beam of thermal neutrons

## Slide 14: Flower behind Al plate Look through metal with thermal neutron beam at NPI Rez by means of Medipix detector with 6LiF converter
Seeds

## Slide 15: Neutron microtomography

## Slide 16: Cold neutron radiography performed at PSI Wrist watch
Notes: 16

## Slide 17: Neutron radiography with Medipix coated by 6LiF: Blank cartridge. Look through metal!
Photography
Neutronogram
Blank shell
(cartridge)
Explosive filling
Roentgenogram

## Slide 18: Test of Spatial Resolution Medipix2 Quad system coated with 6LiF converter illuminated by cold neutrons
100mm
50mm
=>Resolution 65mm!

## Slide 19: 3D-visuallization of proton tracks in silicon pixel detector recorded by Timepix device. Illumination under different angles (0 and 76 degrees) and different applied detector biases (7.2 V and 30 V)
What is the spatial resolution? X- and Y-coordinates are determined with a precision of about 500nm. Determination of angle is with a precision of about 1o. It needs additional experiments.

## Slide 20: Principle of slow neutron detection using 6LiF converter About detection efficiency and position resolution
Simulations performed using MCNP, SRIM and Matlab
Aim: To estimate detection efficiency and spatial resolution
FWHM
PSF
Absorption
UCN
Bragg
curve
6LiF
a
3H-
Silicon
Range of 3H-
Range of a
Insensitive Si layer
Geometry used in simulations
Expected UCNs velocity: 500 cm/s. For such neutrons the cross section of 6Li increases to 0.34 Mbarn. The cross section of 10B reaches 1.67 Mbarn.
50% of such UCNs are fully absorbed in 6LiF layer of 85 ug/cm2 (~ 320 nm thickness). For 10B it is layer of 7 ug/cm2 (~ 30 nm thickness).
Used 6LiF density of 2.65 g/cm3 and 10B density of 2.35 g/cm3.
At 6LiF converter thickness of about 5 mg/cm2, the thermal neutron detection efficiency achieved is about 5%.

## Slide 21: High resolution position sensitive neutron detection based on 10B(n, α)7Li reaction
Pixel detector response to every charged article (alpha and 7Li) in a form of 3D-cluster with a shape corresponding to convolution of Gaussian and individual pixel responses
Corresponding amplitude spectrum measured by integration of cluster volumes. 10B converter thickness 1.8 ug/cm2 (~ 36 nm)
The spectroscopy of alpha particles and 7Li and analysis of their clusters recorded
by Timepix detector permits to achieve spatial resolution up to 3 micrometers

## Slide 22: Neutron images with Timepix detector and resolution calibrated by Siemens star
Alpha clusters centroiding => resolution ~ 3 um
50 µm

## Slide 23: Response of Medipix2 device to natural background radiation This slide inspired the concept of particle tracking and recognition
Clearly recognizable tracks and traces of X-rays, electrons generated mostly by gamma rays, alpha particles, muon, … . Muon tracks can be recognized by submicrometric precision.

## Slide 24: Response of MEDIPIX2 detector with CH2 converter to fast neutrons (17MeV)
The direction of the neutrons with respect to the image was upstream (from bottom to top). The huge background is due to gamma rays which accompany neutrons. Half of the sensor (the right-hand side) was covered with a CH2 foil about 1.3 mm thickness.
One can clearly recognize long and rather thick tracks of recoiled protons (up to 2 mm, vertically oriented) and big tracks and clusters generated via 28Si(n,α)25Mg, 28Si(n,p)28Al nuclear reactions in the body of the silicon detector. These events are displayed on the dense background caused by tracks and traces of electrons from interactions of gamma rays. One can even recognize that proton tracks shapes follows a Bragg law.

## Slide 25: 252Cf neutrons measured in counting mode at different thresholds
252Cf, low threshold (8 keV)
PE
PE + Al
Al
Al
252Cf, high threshold (300 keV)
6LiF
PE
PE+Al

## Slide 26
252Cf, low threshold (8 keV)
PE
PE + Al
Al
Al
252Cf, high threshold (300 keV)
6LiF
PE
PE+Al
252Cf neutrons measured in counting mode at different thresholds

## Slide 27: Responses to different neutron spectra measured at high threshold in counting mode
Thermal neutrons – 500s
252Cf – 2000s
AmBe – 2000s
PE
PE + Al
Al
Uncovered
LiF
17 MeV neutrons at 0o
Identification of spectral composition of neutron radiation fields can be done by comparing responses of different sensitive regions.

## Slide 28: Responses to neutrons of different energies measured at high threshold in counting mode
Thermal neutrons – 500s
252Cf – 2000s
241AmBe – 2000s
PE
PE + Al
Al
Uncovered
LiF
17 MeV neutrons at 0o
Identification of spectral composition of incoming neutron radiation can be done by comparing responses of different sensitive regions.

## Slide 29: On-line radiation monitoring in ATLAS experiment at LHC (16 devices installed within ATLAS)
MPX01 | between ID and JM plug
MPX02 | between ID, LARG and JM
MPX03 | between LARG and LARG EC
MPX04 | between FCAL and JT
MPX05 | between LARG and JT wheel
MPX06 | between LARG and JT wheel
MPX07 | top of TILECAL barrel
MPX08 | top of TILECAL EXT. barrel
MPX09 | corner between JF cyl. and hexagon
MPX10 | cavern wall A or C side
MPX11 | cavern wall USA side
MPX12 | small wheel
MPX13 | between ID and JM plug
MPX14 | between ID, LARG and JM
MPX15 | at the back of Lucid detector
MPX16 | in the USA15 cavern

## Slide 30
All ATLAS-TPX devices were irradiated with thermal (25 meV) and fast (Cf, AmBe) neutrons from well characterized sources in the Czech Metrological Institute.
The reference device was also irradiated with fast 16.7 MeV neutrons from D+T reactions in the IEAP VdG accelerator.
Layout of neutron converters in ATLAS-TPX sandwich device
ATLAS-TPX network
Neutron Efficiency Calibrations of sandwich device

## Slide 31: Fast neutron ToF measurement with TIMEPIX LANSCE neutron sources and nuclear science flight paths (combined ToA and ToT modes)
The layout of the LANSCE neutron sources and Nuclear Science flight paths
Time structure of the proton beam for typical Target-4 operation

## Slide 32
Neutron efficiency calibration from ToA (Los Alamos)
Heavy Tracks and Blobs from 0 to 180 degrees
(0 degree)

## Slide 33
Neutron efficiency calibration from ToA (Los Alamos)
Heavy Tracks and Blobs from 0 to 180 degrees
(180 degree)

## Slide 34: Examples of heavy ionizing events induced by neutrons of different energies selected according their cluster sizes
Different colors and black numbers assigned to individual clusters indicate the energy of incoming neutrons as measured by means of the TOF technique

## Slide 35
SATRAM payload (arrow) onboard ESA Proba-V satellite.
ESA Vega-2 launcher rocket upper stage
ESA Vega-2 rocket

## Slide 36: Timepix/ESA Proba-V: LEO space radiation @ 820 km
LETPs: Energetic light charged particles (I) + nuclear interactions

## Slide 37: Types of neutron detectors already developed at IEAP CTU
The developed detectors are based on hybrid technology connecting a pixelated semiconductor sensor with R/O Timepix chip. The chip (developed within the Medipix2&3 collaboration) contains 65k pixels of 55x55µm2 per 2 cm2. It permits to register a position of an interacting ionizing quantum in the sensor and its ToA and/or ToT.
Presently available sensors: Si, GaAs, CdTe, SiC, 3D-Si, stuffed 3D-Si.
Convertors based on 10B and 6Li are used to detect ultra-cold, cold, thermal and epithermal neutrons, and CH2 for fast neutrons.
Sensitive areas from 2 cm2 to 200 cm2 (with dead area less than 1%).
The fast R/O systems transferring the measured data from different developed Timepix based devices to PC by USB2, USB3, Ethernet interfaces, are available (some of them are displayed on the right).
A prototype of position sensitive detector of neutrons based on crossed strip sensors (resolution of about 30 µm2), what permits to increase neutron detection efficiency by stacking of sensors and place R/O electronics out of neutron beam, has been also developed.

## Slide 38: Basic parameters and fields of application of pixel neutron detector devices developed and tested in IEAP CTU
Position sensitive detection of neutrons with resolution from 3 µm (applying neutron cluster analysis) to 50 µm (in counting mode).
Noiseless neutron detection efficiency up to 7% in case of planar arrangement of neutron converter to the pixelated sensor. Application of "stuffed" 3D-sensor technology promises an increase of the efficiency up to 20%.
Capability to measure time of single neutron arrival with 1.6 ns resolution is well applicable for high resolution ToF experiments.
High-resolution neutron radiography and tomography with ultra-cold, cold, thermal and epithermal neutrons as illustrated above by images of bell flower, a lighter and of the Siemens star of reference.
Energy dependent radiography and tomography performed with resonance neutrons by means of ToF technique.
Simultaneous neutron and X-ray radiography and tomography.
Fast neutron radiography.
Neutron diffraction, elastic and inelastic scattering experiments on relatively small samples using ToF technique.
Measurement of composition and spectral characteristics of mixed neutron-gamma radiation fields and their dosimetry through the network of Timepix wireless detection devices.

## Slide 39: Acknowledgement
The presented results and many others have been achieved within research activities cultivated at IEAP CTU in Prague. They result from extensive partnerships in frame of the Medipix2@3 collaboration with significant contributions of the following colleagues:
R. Ballabriga2, B. Bergmann1, P. Burian1,12, I. Caicedo1, M. Campbell2, J. Dammer1, C. DaVia8, J. Dudák1, C. Froejdh9, E. Froejdh2, V. Georgiev12, C. Granja1, E. Heijne1,2, M. Holík1,12, R. Hall-Wilton10, T. Holý1, J. Jakůbek1, M. Jakůbek1, J. Kirstead7, V. Kraus1,12, F. Krejčí1, E. Lehmann11, C. Leroy4, X. Llopart2, J. M. O‘Donnel3, R. Nelson3, M. Nessi2, A. Owens5, L. Pinsky6, S. Petersson9, S. Pospíšil1, M. Platkevič1, K. Smith13, T. Slavíček1, P. Soukup1, M. Suk1, J. Šolc1, H. Takai7, G. Thungstroem9, D. Tureček1, J. Uher1, D.Vavřík1, Z.Vykydal1, S. Wender7, J. Žemlička1
1 Institute of Experimental and Applied Physics, CTU in Prague, Czech Republic
2 CERN, Switzerland
3 LANSCE, LANL, USA
4 Université de Montréal, Canada
5 ESA
6 NASA/University Houston, USA
7 BNL, USA
8 Manchester University, UK
9 MidSweden University, Sundsvall, Sweden
10 ESS, Sweden
11 PSI, Switzerland
12 WBU Pilsen, Czech Republic
13 Glasgow University

## Slide 40: Now it is the right time for experiments with Timepix detectors managed by Michael Holík and Vladimír Vícha We all are looking forward to meet you sometime in Prague!
Thank you for your attention!