## Slide 1: Waveform Digitizing
Stefan Ritt, Paul Scherrer Institute, Switzerland IEEE NPSS Workshop on Applicatioins of Radiation Instrumentation Nov. 25th, 2022, Dakar, Senegal

## Slide 2: 2 Principles of Detection of Ionizing Radiation
Detectors convert property to be measured directly into electrical signal  position, time
Indirect via light generation in scintillator  energy, time
Silicon detector
Wire chamber
Photomultiplier

## Slide 3: Modern Digital Oscilloscopes
Oscilloscope
Detector
Amplifier
Digitization
Computer
Display
time
01101

## Slide 4: Analog-to-Digital Conversion
ADCs & More

## Slide 5: Digitization: Peak-sensing ADC
-
+
Vin
Vout
C
D
Reset
Amplitude
constant
Technique to digitize fast signals with slow ADCs
Needs reset after each signal
Mainly historical

## Slide 6: Digitization: Charge integration
Charge
Many detectors have a proportionality between charge and particle energy
Noise immunity
-
+
Vin
Vout
Vout
time
Amplitude
Gate

## Slide 7: Digitization: Waveform sampling
time
Amplitude
time
Vth
Tinterp.
Vpk
Qint
Amplitude
ADC Type | Resolution (bits) | Conversion rate
Dual Slope | 12-20 | 100 sample/s
Successive approximation | 8-18 | 10 Msample/s
Flash | 4-12 | 10 Gsample/s
Pipeline | 8-16 | 1 Gsample/s
Delta-sigma | 8-32 | 1 Msample/s
Sample waveform at discrete points (sampling rate) and extract features in digital world

## Slide 8: Digitization: Flash ADC 1-bit and 2-bit
Threshold
VREF
VIN
B
1-Bit Flash
Analog to Digital
Converter
R
R

## Slide 9: Digitization: Flash ADC n-bit
Flash ADC very fast
Requires 2n comparators
Typically <=8 bit resolution

## Slide 10: Digitization: Successive approximation ADC

## Slide 11: Digitization: Pipeline ADC
Combine several flash ADCs with successive approximation logic
Only requires 4-Bit flash ADC
Can convert one sample in each clock cycle
Has a latency depending on the number of pipeline stages
Most common technology for fast ADCs b)
00
01
10
11
Residual voltage
quantized by 2nd ADC
 101110

## Slide 12: ADC Datasheets

## Slide 13: Nyquist-Shannon Sampling Theorem
fsignal < fsampling /2
fsignal > fsampling /2
Only signals with frequencies below half the sampling frequency can be perfectly sampled. Then the original signal can be recovered from the discrete sampling points.

## Slide 15: Limits of waveform digitizing
Aliasing occurs if fsignal > 0.5 * fsampling
Features of the signal can be lost (“pile-up”)
Precise time measurement and good energy resolution need very fast high resolution ADCs

## Slide 16: What are the fastest detectors?
Micro-Channel-Plates (MCP)
  Photomultipliers with thousands of tiny channels (3-10 mm)
  Typical gain of 10,000 per plate
  Very fast rise time down to 70 ps
70 ps rise time  4-5 GHz BW  10 GSPS
SiPMs (Silicon PMTs) are also getting < 100 ps
J. Milnes, J. Howoth, Photek

## Slide 17: Can it be done with FADCs?
8 bits – 3 GS/s – 1.9 W  24 Gbits/s
10 bits – 3 GS/s – 3.6 W  30 Gbits/s
12 bits – 3.6 GS/s – 3.9 W  43.2 Gbits/s
14 bits – 0.4 GS/s – 2.5 W  5.6 Gbits/s
1.8 GHz!
24x1.8 Gbits/s
Requires high-end FPGA
Complex board design
High FPGA power
V1761: 2 Channels, 4 GS/s, 10 bits
Costs: 1-10 k$ / channel
What about 1000+ channels?

## Slide 18: Digitization: Application Specific Integrated Circuit (ASIC) for Waveform Sampling: Switched Capacity Arrays (SCA)
Shift Register
Clock
IN
Out
Waveform
stored
Inverter “Domino” ring chain
0.2-2 ns
FADC
33 MHz
10-100 mW

## Slide 19: Time Stretch Ratio (TSR)
dts
dtd
Typical values:
d ts = 0.5 ns (2 GSPS)
d td = 30 ns (33 MHz)
 TSR = 60

## Slide 20: Triggered Operation
sampling
digitization
lost events
sampling
digitization
Sampling Windows * TSR
Chips usually cannot sample during readout ⇒ “Dead Time”
Technique only works for “events” and “triggers”
Dead time =
Sampling Window ∙ TSR
(e.g. 100 ns ∙ 60 = 6 ms)

## Slide 21: How to measure timing best?
J.-F. Genat et al., arXiv:0810.5590 (2008)
D. Breton et al., NIM A629, 123 (2011)

## Slide 22: How is timing resolution affected
voltage noise Du
timing uncertainty Dt
signal height U
rise time tr
number of samples on slope

## Slide 23: How is timing resolution affected?
U | Du | fs | f3db | Dt
100 mV | 1 mV | 2 GSPS | 300 MHz | ∼10 ps
1 V | 1 mV | 2 GSPS | 300 MHz | 1 ps
1 V | 1 mV | 10 GSPS | 3 GHz | 0.1 ps
today:
optimized SNR:
next generation:
- high frequency noise
- quantization noise

## Slide 24: First Switched Capacitor Arrays
IEEE Transactions on Nuclear Science,
Vol. 35, No. 1, Feb. 1988
50 MSPS in 3.5 mm CMOS process

## Slide 25: Switched Capacitor Arrays for Particle Physics
STRAW3
TARGET
LABRADOR3
AFTER
NECTAR0
SAM
E. Delagnes
D. Breton
CEA Saclay
G. Varner, Univ. of Hawaii
0.25 mm TSMC
Many chips for different projects (Belle, Anita, IceCube …)
0.35 mm AMS
T2K TPC, Antares, Hess2, CTA
H. Frisch et al., Univ. Chicago
PSEC1 - PSEC4
0.13 mm IBM
Large Area Picosecond Photo-Detectors Project (LAPPD)
www.phys.hawaii.edu/~idlab/
matacq.free.fr
psec.uchicago.edu

## Slide 26: DRS4 Chip
Developed at PSI in 2008 together with R. Dinapoli
5 Gsamples/s, 12 bits resolution, 8+1 channels, 17.5 mW/channel
Time measurements down to 10 ps
DRS4 Chip
Evaluation Board
WaveDREAM Board
9000+ channels of MEG II, PSI

## Slide 27: Pulse shape discrimination
g
a
m
g
a
m
Events found and correctly processed 2 years (!) after the data has been taken

## Slide 28: Applications of SCA
Gamma-ray astronomy
Magic
CTA
Antarctic Impulsive Transient Antenna
(ANITA)
320 ps
IceCube
(Antarctica)
Antares
(Mediterranian)
ToF PET (Siemens)

## Slide 29: Conclusions
SCA technology offers tremendous opportunities
Several chips and boards are on the market for evaluation
New series of chips on the horizon might change front-end electronics significantly