## Slide 1
Basics of PET Physics
Margaret E. Daube-Witherspoon, PhD
daubewit@pennmedicine.upenn.edu
January 27, 2023

## Slide 2
Topics of review
Basic nuclear physics as applied to positron emission tomography
Sources of radiotracers for PET
Detection of annihilation photons from positron decay
What is measured with PET
Corrections for physical effects in PET imaging
Limits on PET scanner performance
Time-of-flight PET
Total-body PET imaging considerations & applications
Scanner harmonization

## Slide 3
Nuclear physics
Unstable isotopes
(99mTc – 43 protons, 56 neutrons)
18F – 9 protons, 9 neutrons
11C – 6 protons, 5 neutrons
13N – 7 protons, 6 neutrons
15O – 8 protons, 7 neutrons
Stable isotopes
(98Tc – 43 protons, 55 neutrons)
19F – 9 protons, 10 neutrons
12C – 6 protons, 6 neutrons
14N – 7 protons, 7 neutrons
16O – 8 protons, 8 neutrons
Physical signal

## Slide 4
Gamma decay
99mTc t1/2 = 6 hours, 140 keV
201Tl t1/2 = 3 days, ~70 keV
111In t1/2 = 2.8 days, 171/245 keV
123I t1/2 = 13 hours, 159 keV
Nucleus decays to lower energy state by emitting gamma ray
99Tc (2.13**105 y)
0.0
0.1405
0.1426
99mTc (6.02 h)
g1
g2
g3
Physical signal

## Slide 5
i) Unstable parent nucleus
ii) Proton decays to neutron
Emits positron (e+) and neutrino
vi) Two 511 keV photons
produced by e+ e-
annihilation
~180˚
iii) Positron travels
short distance in tissue
(Neutrino escapes)
Positron decay
11C t1/2 = 20 minutes
13N t1/2 = 10 minutes
15O t1/2 = 2 minutes
18F t1/2 = 110 minutes
Physical signal

## Slide 6
Sources of positron-emitters
Generators
  Useful when radionuclide half-life is short and a radioactive “parent” is available
  68Ga from 68Ge
  82Rb from 82Sr
  Parent radionuclides shipped to hospitals
Cyclotrons
  Necessary when radionuclide half-life is short but NO radioactive parent is available
  18F, 11C, 13N, 15O
  Shorter radionuclides produced on-site
  18F shipped world-wide
Signal source

## Slide 7
Molybdenum generator
99Mo 99mTc
t1/2 = 66 h t1/2 = 6 h
Generators:
Parent-daughter relationships
Transient equilibrium
dNd
dt
= lpNp - ldNd
dNp
dt
= - lpNp
Parent
Daughter
Produced
Decaying
{
{
Signal source
Parent sticks to ion exchange column
Daughter is removed (“eluted”) from the column using chemical differences between parent and daughter

## Slide 8
Generators for PET
dNd
dt
= lpNp - ldNd
dNp
dt
= - lpNp
Parent
Daughter
Produced
Decaying
{
{
Signal source
Gallium generator
68Ge 68Ga
t1/2 = 288 d t1/2 = 68 min
Rubidium generator
82Sr 82Rb
t1/2 = 25 d t1/2 = 75 s
Used in myocardial perfusion imaging
Used to label agents for prostate imaging
Generators must be replaced
Physical decay of parent (less activity available)
Physical decay of column (parent “breakthrough”)

## Slide 9
Cyclotron: isotope production for most
positron emitters
Exponential growth of ‘daughter’
p (22 MeV) + 18O n + 18F
t1/2 ~ 2 h
Signal source

## Slide 10
Cyclotron production
Signal source
Target material
  Robust to irradiation
  Easily handled form (i.e., solid, gas)
Radiochemical separation of radionuclide from target material
Radiochemical production of desired radiotracer
  Adding the radionuclide to a physiologically interesting molecule
  18F in [18F]-fluorodeoxyglucose

## Slide 11
18F-Fluorodeoxyglucose (FDG)
OH
H
CH2OH
OH
OH
OH
Ido et al. 1978
Glucose is key source of energy for many cells, especially tumors
Glucose
Blood  tissue  cell
Phosphorylation … glycogen
O
Signal source
H
H
H
H

## Slide 12
18F-Fluorodeoxyglucose (FDG)
OH
H
CH2OH
OH
OH
H
Ido et al. 1978
Glucose
Blood  tissue  cell
Phosphorylation … glycogen
Deoxyglucose
Blood  tissue
Phosphorylation & then trapped in cell
O
Signal source
H
H
H
H

## Slide 13
18F-Fluorodeoxyglucose (FDG)
OH
H
CH2OH
OH
OH
18F
Patient injected activity: 10 mCi (370 MBq) = 3.7 x 108 decays/s
Tracer kinetics: 6 pico-mole ~ 1 nano-gram
Ido et al. 1978
O
Signal source
H
H
H
H
Glucose
Blood  tissue  cell
Phosphorylation  glycogen
Fluorodeoxyglucose
Blood  tissue
Phosphorylation & then trapped in cell
Information which cells are using glucose & how much

## Slide 14
Some PET Radiotracers and Applications

## Slide 15: Scintillation Detector
Scintillator
Absorbs some or all 511-keV energy & converts to light photons
Scintillator | NaI(Tl) | BGO | LSO/LYSO
t (ns) | 230 | 300 | 40
Photons/MeV | 41,000 | 7000 | 26,000
m (cm) | 2.9 | 1.1 | 1.2
% effic. 2.5-cm | 58% | 91% | ~ 80%
Fast decay (t) -> high count-rate performance good timing resolution
High light output (photons/Mev) -> good energy resolution good spatial resolution (Anger-logic)
High stopping power (m) ->
good efficiency
Signal source

## Slide 16: Photo-sensor
Photomultiplier
Tube (PMT)
converts energy of light photons to electrons (signal)
Position-sensitive (Anger) detector
Block detector
(Anger-logic)
Signal source
PMTs generally larger than small crystals
Requires Anger logic to distinguish crystal of interaction

## Slide 17: Newer Photo-sensor: Si-PMs
Avalanche photodiode (APD)
Small APDs operating in Geiger mode with hundreds of micro cells per mm2:
  Small, compact design
  Available in arrays
  Very high quantum efficiency
  Fast timing characteristics
  High gain, no need for amplification
  Can operate inside MR
  Potential for low encoding (e.g., 1-to-1)
Micro-cell
Common readout line
30 μm
1 mm
Lab. SiPM, Dept. of Experimental and Theoretical Physics, Moscow Engineering Physics Institute
  3.86 x 3.86 x 19 mm3 LYSO
  PDPC digital SiPM
  64-channel array
2D flood
Signal source
Notes: Solid-state photo-sensors introduced in mid-2000’s and used initially for small animal PET scanners
SiPMs have become a real game changer for clinical PET systems in only the last few years
More expensive than conventional PMTs, but the favorable properties make it very desirable for PET, both PET-MR and PET-CT

## Slide 18
Coincidence Detection
coincident photons
c = 3 x 108 m/sec
2t ~ 4-5 ns
Signal detection

## Slide 19
Using coincidence information to create an image
180 views
2 views
4 views
8 views

## Slide 20
  Y = N(T)
What is Measured with PET
a
b
What is measured
Trues
True coincidence
Coincidence?
Record event
Normalization
Signal detection

## Slide 21
  Y = N(Abx T)
What is Measured with PET
a
b
What is measured
Trues
True coincidence
Coincidence?
Record event
Normalization
Attenuation
Signal detection

## Slide 22
  Y = N(Abx T + S)
What is Measured with PET
a
b
What is measured
Trues
True coincidence
Scattered coincidence
Coincidence?
Record event
Normalization
Attenuation
Scatter
Signal detection

## Slide 23
  Y = N(Abx T + S + R)
What is Measured with PET
a
b
What is measured
Trues
True coincidence
Scattered coincidence
Random coincidence
(~2t . singles2)
Coincidence?
Record event
Normalization
Attenuation
Randoms
Scatter
Signal detection

## Slide 24
Interactions of gamma rays with tissue
atom
hno
e-
energy = hno
e-
hn
e-
Photo-electric
Complete absorption of photon by inner-shell atomic electron
Compton
Scattering of photon by outer-shell electron leads to loss of energy and change of direction
Signal interaction

## Slide 25
Attenuation and scatter
d
I = I0 e-md
Narrow beam absorption is exponential with linear attenuation coefficient m
m decreases as photon energy increases
Attenuation decreases detector rates
Broad beam scatter increases detector rates
Signal interaction

## Slide 26
Relative effect of attenuation
SPECT: Lower energy photons compared to PET (larger m), but some photons only travel short distances.
PET: Higher energy photons so  is smaller, but pair of photons must traverse entire body width.
patient
SPECT camera
d1
d2
patient
PET
d
I/I0= e-d1 e-d2 = e-(d1+d2) = 0.06
I/I0 = e-d = 0.20
Suppose d=15 cm at torso center:
m(140keV) = 0.15/cm
m(511keV) = 0.096/cm
Suppose d1=d2=15 cm:

## Slide 27
In SPECT, we do not know d; therefore, only approximate correction is possible
Attenuation correction is more accurate in PET than SPECT
In PET, the total path length D=d1+d2 can be measured and allows an accurate correction
patient
SPECT camera
d1
d2
patient
PET
d?
I/I0 = e-d?
I/I0 = e -d1 e-d2 = e-D

## Slide 28
Transmission Scan
I / I0 = e-md1 . e-md2
= e-mD
Emission
137Cs point source (662 keV, t1/2 = 30 yr)
68Ge line source (511 keV, t1/2 = 288 d)
d1
d2
Transmission
I / I0 = e-mD  AC = emD
D

## Slide 29
X-ray CT has replaced gamma-ray transmission sources for attenuation measurements
X-ray CT scan - source of X-rays with energies from ~30 to 120 keV. We assume an ‘effective’ energy of ~ 75 keV
E (keV)
30
120
511
662
Intensity
I0(E)
X-ray source
positron source
g-ray source
0
spectra
(Recall that the PET emission data is attenuated at 511 keV)
511 keV
CT -> PET Transmission image: scaled and hybrid segmentation

## Slide 30
Impact of Attenuation Correction
No AC
AC
Without attenuation correction:
Skin artifact
Artificially high lung uptake
Lower apparent uptake in the center

## Slide 31: Randoms Correction — Delayed window technique
A
B
Delayed Signal A
Signal B
Coincidence
Window, 2
Signal A
time
Randoms = (2 x singles**2
True coincidence
Random coincidence

## Slide 32: Randoms Correction — Calculation from singles rates
Randoms = (2 x singles**2
Measured randoms depends on the rate of measuring coincident events  noisy correction
Coincidence fraction (# coincidences / # singles) is a few %
Calculation of randoms from singles depends on the rate of measured single events  less noisy correction
Noise in corrections can propagate into image
Maps of randoms:
Measured
Calculated

## Slide 33: Scatter Correction
A
B
Single Scatter Simulation:
Calculate the contribution to each LOR for an arbitrary scatter point using the Klein-Nishina equation for single Compton scatter with known attenuation distribution + guess of activity distribution
before
after

## Slide 34
True events
Scatter events
threshold with poor energy resolution -> more scatter events measured
scatter
random
true
3D: higher sensitivity…increased scatter
threshold with good energy resolution -> fewer scatter events measured
Energy rejection important for 3D
2D septa allowed mechanical rejection of scatter, randoms

## Slide 35: Scanner Deadtime
NEMA NU2-2001
20-cm phantom
Philips Allegro
Univ. of Pennsylvania
NEC = T/(1+S/T+R/T)
Ideal Trues
(no deadtimes)

## Slide 36
Clinical imaging range
20 cm
27 cm
35 cm
diam.
Image quality (SNR) depends on patient size
NEC ~ (Signal / Noise)2
Signal/noise ratio
Activity
Notes: These data are from Allegro, but I would make argument general - heavy patients have lower image quality, and it is not practical to make up the difference in ’good’ counts (~6x) by imaging longer, or by injecting more activity.

## Slide 37: PET Scan Quantification Standardized Uptake Value (SUV)
What is it?
  SUV = (Activity Concentration in region) / (Activity Concentration in whole volume)
Why use SUV?
  Repeat scans with different injected activities
  Repeat scans on scanners with different sensitivities
  Clinical trials – conformity of results among sites
Conc = Dose/Weight
ALESION = 7.8 kBq/mL

## Slide 38: PET Scan Quantification Standardized Uptake Value (SUV)
What is it?
  SUV = (Activity Concentration in region) / (Activity Concentration in whole volume)
Calibrated from uniform phantom study
Conc = Dose/Weight
ALESION = 7.8 kBq/mL
SUVLESION = 1.8

## Slide 39
FDG PET/CT - staging lung cancer
Lung
Pelvis
August 9
October 10
What is the quantitative change in lesion uptake?

## Slide 40
Chain of data quantitation
Accuracy depends on many factors: dose assay, instrument calibration, reconstruction
SUV depends on lesion size and image reconstruction approach, e.g. filter and #iterations
Courtesy, Siemens

## Slide 41
Limits on PET scanner performance
Limitations of scanner
  Spatial resolution
  Sensitivity (geometric scanner design, detector efficiency)
  Count rate performance (non-linear behavior as increase activity)
  Inaccurate corrections for physical effects
Limitations by patient
  Motion
  Tracer uptake

## Slide 42: Limits on spatial resolution
R
Positron range, R:
 | 18F | 11C | 82Rb
Rmax (mm) | 2.6 | 3.8 | 16.5
FWHMp (mm) | 0.22 | 0.28 | 2.6
~180˚
Photon non-collinearity:
  FWHMNC=0.0022 X scanner diameter
  (2mm for a 90-cm diameter)
Detector resolution (FWHMd )
  Crystal size

## Slide 43: Impact of spatial resolution limitations
Spatial resolution blurring impacts
Quantitative accuracy (blurring of uptake between structures)
Detection of small tumors / structures
Depends on location in scanner FOV
Partial volume effect (PVE)
Measured activity < true activity
Recovery coefficient = Ameas / Atrue
SUVmean – average uptake in volume of interest (VOI) over structure
SUVmax – maximum uptake in structure of interest (1 voxel)
SUVpeak – average uptake in small VOI within structure

## Slide 44: Time-of-flight (TOF) PET
Image reconstruction from projections
back-projection

y(s,)
x(u,v)
Low noise
Low contrast
High noise
High contrast
non-TOF
Low noise
High contrast
TOF
TOF information reduces coupling of signals, thus improves SNR
x = c . t/2
500 ps  ~7.5 cm
220 ps  ~3.3 cm
10 ps  1.5 mm
t = t1-t2
t1
t2
TOF kernel
Notes: One of the key concepts originally introduced in the late 1970’s and refined in the last several years is TOF. In brief - the precise measurement of the time difference between detectors allows the reconstruction to achieve higher SNR, approximately a factor of 2x for current instruments.
We see here the benefit – without TOF we can achieve high contrast – better quantitative accuracy – with a larger number of iterations which also leads to higher noise. With TOF the faster local convergence allows us to achieve both high contrast and low noise – this is often interpreted as an effective gain in sensitivity.

## Slide 45: Philips Ingenuity TF
Siemens mCT
GE Discovery 690, 710
TOF PET/CT scanners from all vendors
PMT-based PET/CT: 2006 -> Timing resolution 500-600 ps
SiPM-based PET/CT: 2017 -> Timing resolution 200-400 ps
GE Discovery MI
Philips Vereos
Siemens Vision

## Slide 46
What is TB (Total-Body) PET?
TB-PET: Higher sensitivity per volume
  Total sensitivity gain (>20x for 1-m axial FOV)
  Per-organ sensitivity ~3x
Simultaneous imaging of large volume
  Kinetic modeling of multi-organ system
Commercial PET-CT: Axial FOV 15-30 cm
Notes: Gain in total-body sensitivity is very high, > 20x for 1-m scanner
Gain per organ (point) is more modest, about 3x

## Slide 47
Long axial FOV scanners
uEXPLORER PET Detectors: 10:1 encoding, crystals:SiPMs
Block with integrated light-guide
5x14 blocks
7x6 crystals/block
LYSO Detector Module
18 mm
2.76 mm
6 mm
2940 crystals
Block flood map
uEXPLORER (United Imaging)
Courtesy, Hongdi Li, UIH America
TOF = 509 ps, 3-mm spatial resolution
PennPET Explorer
TOF = 250 ps, 4-mm spatial resolution
PennPET Detectors: 1:1 encoding, crystals:SiPMs
  3.86 x 3.86 x 19 mm3 LYSO
  PDPC digital SiPM
2D flood
194 cm AFOV
64 - 142 cm AFOV
Notes: Leading into the next talk – what do we need for long AFOV scanners where cost will be an issue
Both uEXPLORER and PennPET Explorer have chosen to use LYSO but in very different detector configurations. UI has better spatial resolution, PennPET has better TOF.

## Slide 48
Long axial FOV scanners
Siemens Biograph Vision Quadra
Prenosil, et al., J Nucl Med, 2021
TOF = 230 ps, 3.5-mm spatial resolution
106 cm AFOV
Quadra Detectors: 25:1 encoding, crystals:SiPMs
GE Omni Legend
BGO detectors with deep learning to obtain TOF information
32-64 cm ( 128 cm) AFOV
Notes: Leading into the next talk – what do we need for long AFOV scanners where cost will be an issue
Both uEXPLORER and PennPET Explorer have chosen to use LYSO but in very different detector configurations. UI has better spatial resolution, PennPET has better TOF.

## Slide 49: How do we measure benefit of new scanner performance or image reconstruction approach?
Lesion contrast (SUV) - accuracy of quantitation
Detectability - numerical and human observers
Variability of lesion uptake - as a function of location, environment, statistics
Must define task:

## Slide 50
Tracer: e.g., FDG biology and time of uptake
Physics: b+ range, b+ fraction
Patient variables: motion (breathing) habitus (BMI)
SNR: scan time instrument performance - sensitivity, spatial resolution, TOF
Quantification: dose and instrument calibration data correction methods reconstruction algorithm
Tracer uptake measurements depend on…

## Slide 51: Physical phantoms to assess performance of image quality
NEMA Image Quality
SNMMI CTN

## Slide 52
Standardizing Performance with NEMA IQ Phantom
8.5 mm
44 mm
Default parameters
Harmonized parameters
5 scanners at 3 sites
Notes: To better understand what we are trying to do, I’ll first show the data we are collecting.
The data collected for this project was collected with the NEMA IQ phantom, shown above. It holds six spheres at a time, which are selected from the 13 spheres sizes offered, ranging from 8.5mm to 44mm in diameter.
3. Data is collected for these experiments with hot spheres in a warm background, with two possible contrast ratios: 9.7:1 or 4.85:1

## Slide 53
Quantifying Performance with NEMA IQ Phantom
CRC: Mean VOI
CRC: Max voxel
  Resolution (PSF) model can be included to compensate for parallax errors
Notes: To better understand what we are trying to do, I’ll first show the data we are collecting.
The data collected for this project was collected with the NEMA IQ phantom, shown above. It holds six spheres at a time, which are selected from the 13 spheres sizes offered, ranging from 8.5mm to 44mm in diameter.
3. Data is collected for these experiments with hot spheres in a warm background, with two possible contrast ratios: 9.7:1 or 4.85:1

## Slide 54
Summary
Basic nuclear physics as applied to positron emission tomography
Sources of radiotracers for PET
Detection of annihilation photons from positron decay
What is measured with PET
Corrections for physical effects in PET imaging
Limits on PET scanner performance
Time-of-flight PET
Total-body PET imaging considerations & applications
Scanner harmonization

## Slide 55
What determines the direction of measured events in PET?
  Septa
  Collimator
  Coincidence line between 2 photons detected within timing window
  Choice of scintillator
Notes: Remember, we assume two photons are related if they arrive at detector within a small timing interval.

## Slide 56
Which two factors primarily determine the spatial resolution achieved in a clinical PET scanner during 18F-FDG imaging?
  Positron range
  Detector spatial resolution
  Photon non-collinearity
  Detector timing resolution

## Slide 57: This PET image:
Has not been corrected for attenuation.
Has been over-corrected for attenuation.
Has been corrected for attenuation using a Ge-68 transmission source rather than a CT scanner.
Shows pathologic uptake of the PET tracer in the skin.
Has not been properly normalized.
Notes: Other tell-tale sign of NAC = hot lungs.