## Slide 1: Refresher: Physics of Hybrid Imaging IEEE NPSS School of Advanced Topics in PET/CT and PET/MR
Youngho Seo, PhD
UCSF Physics Research Laboratory – University of California, San Francisco and Berkeley, and Lawrence Berkeley National Laboratory

## Slide 2: Combined SPECT/CT and PET/CT/MRI Developments
The first multimodality prototypes developed in the mid 1990s: (a) a SPECT/CT scanner at UCSF, (b) a PET/CT scanner evaluated clinically at the University of Pittsburgh and (c) an MR/PET scanner for small animals at UC, Davis.
Townsend DW. Phys Med Biol. 2008;53:R1-R39

## Slide 3: Dual-Modality Imaging – Example of SPECT/CT
Functional-anatomical correlation
  Rigid transformation of SPECT into CT
CT-based attenuation correction and scatter correction
  Generate attenuation map from CT, not a source-based transmission scan.
  With CT-derived attenuation map and EM (estimation maximization) algorithm, attenuation correction is ‘standard’ with current SPECT/CT and PET/CT scanners.
CT
CT + SPECT
SPECT (FBP)
MLEM (AC, Geometric Blur Corrected)

## Slide 4
SPECT/CT (or SPECT-CT)
CT/SPECT (or CT-SPECT)
PET/CT (or PET-CT)
CT/PET (or CT-PET)
SPECT/MR(I)
MR(I)/SPECT
PET/MR(I)
MR(I)/PET
Hybrid Imaging
Dual-modality
Combined dual-modality
…

## Slide 5: Software and Hardware Fusion
Software-based image registration is complementary, not competitive with dual-modality imaging
Real-time ultrasound/gated SPECT
(Shekar et al., Radiographics 23:1673, 2003)
SPECT/MRI in basal ganglia with
123I-b-carbomethoxyiodophenyl tropane
(Slomka, et al., J Nucl Med 45:36S, 2004)
SPECT/CT with 111In-ProstaScint®
(Wong, et al., PhD thesis, UCSF, 2002)
Pre/postsurgical PET/CT of head-neck squamous cell carcinoma
(Towsend et al., Semin Nucl Med 33:193)
Notes: SPECT Ultrasound an MRI SPECT images from Shekar et al., Radiographics 23:1673, 2003
Since the beginning of nuclear medicine, users have recognized that radionuclide imaging lacks anatomical information and offers relatively poor spatial resolution. For this reason, there always has bee a strong interest in combining SPECT and PET images with anatomical information from different modalities. One way to do this is to acquire the anatomical and functional information on different systems in separate procedures, and then combine the resulting images in software. This offers the flexibility of allowing the combination of images that don’t exist in a combined system. It can work well in a localized region or where body parts can be constrained so that they don’t move excessively, for example in the head and brain.

## Slide 6: Voluntary and Involuntary Patient Motion
Images courtesy: GE Medical Systems
SPECT/CT
CT only
SPECT/CT images acquired with shallow breathing
Inhalation /Expiration CT Subtraction

## Slide 7: SPECT and Transmission Imaging
Patient
SPECT
Detector
Scanning Radioactive Line Source
Bailey, “Transmission Scanning in Emission Tomography”, Eur J Nucl Med, 25:774-787, 1998.

## Slide 8: Emission-Transmission Imaging Conceived at Beginning of Radionuclide Imaging
Mark II
Univ of Pennsylvania
David Kuhl, MD
On May 14, 1965, Kuhl, Hale, and Eaton obtained the first transverse axial transmission CT scan of a patient’s thorax using a Mark II scanner. The CT scan was accomplished by replacing the detector underneath the patient with a small radioactive source (Am-241) such that the movement of the source underneath the patient followed the movement of the detector above. Kuhl DE, Hale J, Eaton WL. Transmission scanning: a useful adjunct to conventional emission scanning for accurately keying isotope deposition to radiographic anatomy. Radiology 1966:87:278-284.

## Slide 9: Transmission-Emission Computer Tomograph (1987)
DM Mirshanov, Tashkent Branch, All-Union Research Surgery Center, USSR Academy of Medical Science. “Transmission-Emission Computer Tomograph”, USSR Patent 621.386:616-073 20.01.87-SU-181935, Jan 20, 1987.
Semiconductor Detector
Strip Scintillator
X-Ray Source
Display
Image
Memory Buffer

## Slide 10: “Transmission/Emission Registered Imaging (TERI) Computed Tomography Scanners”
Abstract: “Acquisition of CT trans-mission data and emission data is achieved with the same scanner. Because the transmission and emission data sets are spatially registered, the map of attenuation coefficients described by the recon-structed transmission images can be used for generating a mathe-matically correct matrix for correc-tion of sequentially or simultaneous-ly acquired SPECT imaging data. The imaging devices can be gam-ma camera.”
Filing Date: 2 July 1990
Applicant and Inventor: Charles H. Kaplan, Chicago IL
International Patent Application PCT/US90/03722

## Slide 11
SPIE Vol 1231 Medical Imaging IV Image Formation (1990)
Description of a Simultaneous Emission-Transmission CT System
Bruce H. Hasegawa, Eric L. Gingold, Susan M. Reilly,
Soo-Chin Liew, and Christopher Cann
Department of Radiology and Bioengineering Graduate Group
University of California, San Francisco (Box 0628)
San Francisco, CA 94143
1. ABSTRACT
We are designing an instrument which will perform correlated emission-transmission image acquisition, but which departs from previous systems by incorporating a low-power x-ray tube and generator, rather than a radionuclide source, for the transmission image. The system uses an array of high-purity germanium (HPGe) detectors and detector electronics with energy discrimination circuitry to separate x-rays (at 100 or 120 kVp) from higher energy gamma rays from the 99mTc or 123I radiopharmaceutical injected into the patient. The data acquisition electronics have time constants matching the charge collection time (50 ns) of the HPGe detectors to maximize count-rate capabilities (up to 1 million cps per detector element), while maintaining adequate energy resolution (approximately 10% FWHM). Each detector channel has two energy windows for simultaneous transmission-transmission

## Slide 12: Simultaneous SPECT/CT (1990)
Simultaneous Emission (SPECT) – Transmission (x-ray CT) System
  Shortcomings: High x-ray flux compared to gamma emission, unnecessary motion artifacts for x-ray tomographic acquisition, etc.
X-ray tube
collimator
Detector (e.g., HPGe)
Radionuclide source (e.g., 99mTc)
X-ray
Energy Discrimination
X-ray data
Radionuclide data
LN2 cryostat

## Slide 13: Transmission Imaging (Radionuclide vs. X-Ray Source)
1 Ci = 3.7×1010 disintegrations/sec
Assume 3.7×1010 photons/sec
1 Ci source produces
2.94×103 photons/sec/mm2 at 1 m
Siemens Somatom
125 kV, 2.5 mm Al + 0.4 mm Cu, 24º W target
3.5×106 photons/mAs/mm2 @ 0.75 m
R Birch, M Marshall, GM Ardran: Catalogue of Spectral Data for Diagnostic X-rays, Scientific Report Series No. 30 (The Hospital Physicists' Association, London, 1979).
1 mA x-ray tube produces
1.97×106 photons/sec/mm2 at 1 m
CT scanners operate at 600 mA
1.18×109 photons/sec/mm2 at 1 m
1 mA tube equivalent to 669 Ci
600 mA tube equivalent to 401,190 Ci
X-ray source offers faster scans, higher statistics, better spatial resolution, and can be turned off and thrown away

## Slide 14: Is Simultaneous SPECT CT Feasible With a Single Detector?
Filing Date: 2 July 1990
Applicant and Inventor: Charles H. Kaplan, Chicago IL
1 mA x-ray source produces 2 million photons/sec/mm2 at 1 m
Radionuclide image < 1 photons/sec/mm2
Really fast camera count-rate necessary
Possible damage to photomultiplier tubes
Can discriminate x-ray from γ-ray photons with energy-resolving x-ray detectors
Maybe

## Slide 15: Sequential Dual-Modality SPECT/CT
X-ray tube produces photon fluence equivalent to several thousand Curies of radioactivity
X-ray scatter ~ 0.1% primary
X-ray scatter to radionuclide detector can be equivalent of hundreds to thousands of Curies of radioactivity
Separating radionuclide detectors from x-ray beam reduces scatter cross-talk
Camera
Camera
X-ray tube
X-ray detector

## Slide 16: UCSF SPECT/CT System (1995)
Computer
Correlated Image
CT
X-Ray CT
Radionuclide
Emission
SPECT
Equipment courtesy of GE Medical Systems

## Slide 17: Modern Commercial SPECT/CT Systems
GE Discovery NM/CT 670
Siemens Symbia
Spectrum Dynamics Veriton-CT
GE StarGuide

## Slide 18: SPECT-CT Movie
Movie courtesy of GE Healthcare
Disclaimer: Claims embedded in this movie are the views of GE Healthcare.

## Slide 19: Correction of Radionuclide Data using CT
CT Image
Attenuation Coefficients
Object Geometry
Scatter Coefficients
Correct Radionuclide Data for Physical Perturbations
Improve Image Quality and Quantitative Accuracy of Radionuclide Data

## Slide 20: CT-based Attenuation Correction
X-ray CT offers rapid attenuation map generation capability.
Townsend DW, et al. J Nucl Med. 2004;45:4S-14S.
transmission
emission
transmission
emission
transmission
emission
bed 1
bed 2
bed n
CT
emission
emission
emission
bed 1, 2, …
bed n
Transmission-Emission
CT Transmission-Emission

## Slide 21: CT-Based Attenuation Coefficients Derivation
Calibration Procedure
Scan calibration phantom with K2HPO4 on x-ray CT scanner
Reconstruct x-ray CT data
Define VOI’s to get CT numbers (HU’s)
Calculate expected m(EE)
sample chamber
acrylic
water
CT Image of Calibration Phantom
0.00
0.05
0.10
0.15
0.20
Air
60% ETOH
Water
0.9% Saline
100 mg/cc
200 mg/cc
300 mg/cc
400 mg/cc
50 mg/cc
-1000
-800
-600
-400
-200
0
200
400
600
Hounsfield Units
Linear Attenuation
Coefficients (cm )
-1
Piecewise linear fit

## Slide 22: Attenuation Correction Algorithm
Energy
Calibration
Attenuation
Map
Attenuation map must be calibrated for energy of radionuclide photons!
Emission
Projection
Iterative Reconstruction
Transmission
Projection
Filtered
Backprojection
Transmission
Tomogram
Attenuation
Corrected
SPECT/PET
Image

## Slide 23: 131I-MIBG Uptake in Neuroblastoma
CT
SPECT
SPECT/CT
CONVENTIONAL

## Slide 24: Combined PET/CT (1998)
University of Pittsburgh Medical Center (UPMC) in collaboration with CTI (Knoxville, TN) developed:
  Combined CT (Siemens Somatom AR.SP) and PET (ECAT ART, CPS Innovations) – Partial ring PET needed rotation.
Townsend DW, et al. J Nucl Med. 2004;45:4S-14S.

## Slide 25: Commercial PET/CT Development
CPS Innovations (Knoxville, TN) introduced the commercial version of PET/CT.
Townsend DW, et al. J Nucl Med. 2004;45:4S-14S.

## Slide 26: Modern Commercial PET/CT Systems

## Slide 27: Combined PET/MR and SPECT/MR
Using MRI as anatomical reference for PET
Nonionizing radiation.
Excellent soft tissue contrast by MRI.
A means of imaging functional(PET)-functional(e.g., MRS, fMRI)-anatomical data (e.g., T1w, T2w MRI) as one-stop shopping or even simultaneous imaging possible (almost).
  MR-compatible microPET insert
  Human PET head insert
  Whole-body PET/MR as a single gantry or two separate systems with a patient bed that shuttles between the two scanners
Challenges include:
  Accurate attenuation map generation and correction.
  Synergy – Motion correction for PET data using MR data.
Same principles are applied to SPECT/MR (still at a preclinical stage).

## Slide 28: Soft Tissue Contrast of CT and MRI
CT
MRI
Ahmed M, et al. Semin Vasc Surg. 2004;17:181-205 (top images) ,
The Prostate Centre (www.prostatecentre.ca) (bottom left) &
Image courtesy of J. Kurhanewicz (UCSF) (bottom right)

## Slide 29: MR-Compatible microPET Insert
One such example is from the group at University of California, Davis. They developed:
  MR-compatible (for both PET and MRI) small animal PET insert that provides simultaneous PET/MRI capability in a high-field small animal MRI scanner.
Cherry SR. J Nucl Med. 2006;47:1735-1745.

## Slide 30: Total-Body PET/MRI
Ideas to realize total-body PET/MRI include
  Limited axial field of view PET and MRI combination
  Full cylinder of PET and MRI combined
Cherry SR. J Nucl Med. 2006;47:1735-1745.
Total-Body (Full Cylinder) PET
By UC Davis and United Imaging Healthcare

## Slide 31: Total-Body PET/CT
uEXPLORER
Ring diameter: 78.6 cm; Transaxial FOV: 68.6 cm; Axial FOV: 194.8 cm
# of crystals: 564,480; # crystal blocks: 13,440; # of SiPMs: 53,760
80 detector row CT

## Slide 32
Notes: https://www.youtube.com/watch?time_continue=6&v=JaszDkmgfMY

## Slide 33: Integrated PET/MRI Systems
PET Insert
These images(top left, top right, and bottom right) are property of Siemens Healthcare. Scrapped from press releases.
Pichler BJ. J Nucl Med. 2010;51:333-336.
PET Insert
GE Confidential and Proprietary
Technology in development that represents ongoing research and development efforts. These technologies are not products and may never become products. Not for sale. Not cleared or approved by the FDA for commercial availability.

## Slide 34: two is now one. The first whole-body MR-friendly PET architecture

## Slide 35
Magnet
PET Ring
Front bridge
PET/MR Cradle
PET/MR Mobile table
Stationary CMA coil
LPCA
Rear bridge
Rear pedestal
PET Tube
2379mm
1720mm
1600mm
Front end bell**
250mm
PET detector
Rear end bell **
** Other Magnet enclosures are not shown for clarity
GE Time-of-Flight PET/MR

## Slide 36: Two Systems Sharing One Patient Bed
This PET/MRI system from a commercial vendor put together separate PET and MRI systems with one patient bed that goes in both scanners without having to move the patient.
This image is property of Philips Healthcare. Scrapped from a press release.
Zaidi H, et al. Phys Med Biol. 2011;56:3091-3106

## Slide 37: Attenuation Correction of PET using MRI
MRI does not provide linear attenuation coefficient map that is easy to derive from x-ray CT images.
  Use of an MR image for attenuation map generation is works-in-progress.
  Segmentation and atlas matching method, and CT-like pulse sequence that delineates bones are under investigation.
Hofmann M, et al. J Nucl Med. 2008;49:1875-1883.
Hofmann M, et al. Eur J Nucl Med Mol Imaging. 2008;36:93-104

## Slide 38: CT-like MRI (UTE/ZTE) & Motion Correction of PET using MRI
Navigator echo technique from MRI could provide real-time motion-tracking data (including respiratory-induced internal motions).
Simultaneous PET/MRI thus can utilize the periodic MRI navigator signal to correct for motion-induced projection data inconsistency of PET.
Wiesinger et al., ZTE Bone Imaging, ISMRM 2014
Ultrashort or zero echo time (UTE/ZTE) MR pulse sequence provides intense signal from bone, which can create CT-like MR images. Then, potentially, UTE/ZTE MR image can be better utilized for PET attenuation correction.
CTAC (reference)
Zero TE
ZTAC
Yang J, et al. SNMMI. 2016

## Slide 39
Deep Learning Approach Example for Direct AC and SC
Conventional method:
  Pseudo CT generation
  μ-map transformation
  Scatter simulation
  Separately corrected in sinogram space
Deep Joint Correction:
  Accelerated one-step
  Not using CT or MRI
Yang J, et al. Phys Med Biol. 2019;64:0750019
Notes: However, this approach can correct attenuation and scatter simultaneously in image space, without using CT or MR images.

## Slide 40: Positron Range in a Magnetic Field
Cherry SR. J Nucl Med. 2006;47:1735-1745
Then, what happens if these positron emissions occur in a strong magnetic field?
Levin CS, Hoffman EJ. Phys Med Biol. 1999;44:781-799
Radionuclide | Half-life | Max. + energy
11C | 20.4 min | 0.96 MeV
13N | 9.97 min | 1.19 MeV
15O | 123 sec | 1.73 MeV
18F | 109.8 min | 0.634 MeV
68Ga | 67.8 min | 1.90, 0.821 MeV
82Rb | 76 sec | 2.60, 3.38 MeV
89Zr | 78.4 hr | 0.902 MeV
124I | 4.18 days | 1.54, 2.14 MeV

## Slide 41: Positron Range in a Magnetic Field
Left: The simulated positron trajectories of a 68Ga point source in water is shown with a uniform magnetic field at (a) 0 T, (b) 3 T, and (c) 7 T. The magnetic field was applied in the direction going into the paper plane
Right: The simulated positron PSF of 68Ga in water with a 3 T uniform magnetic field
Huang S-Y, et al. IEEE Nuc Sci Symp Conf Rec. 2014

## Slide 42: Positron Range in a Magnetic Field
The system spatial resolution in full-width at half maximum (FWHM) is compared between the PET/CT (orange) and PET/MRI (blue) measured with the (a) 18F and (b) 68Ga line source phantoms imaged in water.
Huang S-Y, et al. IEEE Nuc Sci Symp Conf Rec. 2014
Radionuclide | Half-life | Max. + energy
18F | 109.8 min | 0.634 MeV
68Ga | 67.8 min | 1.90, 0.821 MeV

## Slide 43: SPECT/MR
Although SPECT/CT was developed prior to PET/CT, development of combining SPECT with MRI is slowly progressing. It is the same principle. MR-compatible SPECT (still in the preclinical stage) is being evaluated.
  Replace photomultiplier tubes with solid-state detectors (CdZnTe or CdTe) or APD-based technologies.
  Nonmagnetic collimators (casting of collimators using tungsten powder composite)
  Rotation can be avoided by building a ring detector system for SPECT.
Xu J, et al. Proc of SPIE. 2010;7622:76220V
Lee KS, et al. Phys Med Biol. 2011;56:685-702.