## Slide 1
Virtual Monoenergetic Imaging from Dual sources Dual Energy CT
Presenter: Pham Nguyen Kim Ngan
Ho Chi Minh City, 2025
Department of Medical Imaging and Radiological Sciences, Chang Gung University, Taoyuan, Taiwan.

## Slide 2
2
Fig 2[1]. These artifacts between hip replacements (1a, left), and dental implants made of high molecular weight metals. (1b, right)
Computed Tomography (CT)
Detectors
X-rays source


Pixels
Voxels
Reconstruction
Diagnosis
Oncology
Fig 1. The imaging progress of CT
[1] Wellenberg, R. H.H., E. T. Hakvoort, C. H. Slump, M. F. Boomsma, M. Maas, and G. J. Streekstra. “Metal Artifact Reduction Techniques in Musculoskeletal CT-Imaging.” European Journal of Radiology 107 (October 1, 2018): 60–69. https://doi.org/10.1016/j.ejrad.2018.08.010.
. Introduction

## Slide 3
3
Fig 3[2]. CT number of materials representing Single-energy CT and Dual-energy CT.
Virtual Monoenergetic Image (VMI)
Abnormal delineation
Diagnosis
Metal artefact reduction
40 keV
40-60 keV
>100 keV
Dual-energy Computed Tomography (DECT)
At any arbitrary energy (keV):
Dual-energy CT
Low E spectrum
High E spectrum
Virtual Monoenergetic Images
Material Decomposition
Effective Atomic Number
Electron Density
. Introduction
[2] F. Tatsugami et al. “Dual-energy CT: minimal essentials for radiologists,” Japanese Journal of Radiology, vol. 40, no. 6, pp. 547–559, Jun. 2022.

## Slide 4
. Materials and methods
Numerical simulation data[3]: Mayo Clinic Database: 120 kV, 200 mAs.
4
Patient CT images collection
Clinical data : Gammex RMI 467 phantom: masks and CT images (80kV and 140kV)
80 kV
140 kV
Ideal Mono Image
Mask
80/140 kV
Siemens VMI
[3] Y. Zhang and H. Yu, “Convolutional Neural Network Based Metal Artifact Reduction in X-Ray Computed Tomography,” IEEE Transactions on Medical Imaging, vol. 37, no. 6, pp. 1370–1381, Jun. 2018
+ Changhwa Christian Hospital, Taiwan.
+ Siemen Healthineers Force – Dual source (80/140kVp & VMI:40-190 keV)

## Slide 5
. Materials and methods
k =1,2: Two basic materials
L,H: Low and high energy
Weighting factors for simulation[5][6]:
Fig 4. Attenuation coefficient map of two basic material simulated phantoms at [80;140] kV
[4] L. Yu, S. Leng, and C. H. McCollough, “Dual-energy CT-based monochromatic imaging.,” AJR. American journal of roentgenology, vol. 199, no. 5 Suppl, 2012, doi: 10.2214/ajr.12.9121.
[5] L. Yu, J. A. Christner, S. Leng, J. Wang, J. G. Fletcher, and C. H. McCollough, “Virtual monochromatic imaging in dual-source dual-energy CT: Radiation dose and image quality,” Medical Physics, vol. 38, no. 12, pp. 6371–6379, 2011

## Slide 6
=[kVp]*[w]
w
[kVpclin]\ [VMIclin]
[kVpclin]\ w
clin: clinical CT images
kVpclin: [80,140] kVp
VMIclin: [40:10:190] keV
ID
1 --> LN-300 Lung
2 --> LN-450 Lung
3 --> AP6 Adipose
4 --> Breast
5 --> True Water
6 --> Solid Water
7 --> Brain
8 --> Liver
9 --> B-200 Bone
--> CB2-30%
11 --> CB2-50%
12 --> Cortical Bone
13 --> Inner Bone
Fig 5. Gammex RMI467 phantom: (a) Real image, (b) Index image
. Materials and methods
Weighting factors for clinical data:

## Slide 7
Numerical simulation data
Fig 6: Patient with metal-free slice in the abdomen, comparing between VMIs and Ideal monoenergetic images at: a) 60keV, b) 90 keV and c)120 keV.
III. Results & Discussion
Ideal Mono Image
VMI
Relative Difference

## Slide 8
“An accurate choice of the energy level of VMI reconstruction should take into account different parameters, including the purpose of reconstructing VMI (e.g. increase vascular contrast vs reducing beam-hardening artefacts), but also patient size and optimization of the image quality” [3]
“The range between 140 to 200 keV, artifact reduction reached a plateau. Thus, it might make sense to use 140 keV as a standard reconstruction when aiming to yield a higher soft tissue contrast for the evaluation of pelvic organs. This is in line with two phantom studies examining VMI derived from SDCT and dual source technology proposing 130–150 keV as a possible standard setting for VMI reconstructions in THR case”[4]
Article review
Lateral
Medial
Pubic bone
Muscle
Fat
[5] Tommaso D’Angelo, et al (2019), Dual energy computed tomography virtual monoenergetic imaging: technique and clinical, Br J Radiol 92: 20180546
[6] Kai Roman Laukamp, et al (2018), CT metal artifact in patients with total hip replacements: for artifacr reduction monoenergetic reconstructions and post-processing algorithms are both efficient, but not similar , European Radiology 28:4524-4533
Image quality index
8

## Slide 9
Fig 7: Weighting factors derived from simulation database and clinical database.
Fig 8: The relative difference of Gammex RMI467 phantom VMIs derived from estimated weighting factors and commercial Siemens VMIs.
Clinical data
III. Results & Discussion

## Slide 10
10
Low E images
High E images
Metal Artifact Resistant VMI
DECT images w/wo metal artifacts
PINN
Physic-informed neural network (PINN)[8],[9]
PINN integrate existing knowledge into deep learning models
Model performance is superior to traditional models.
More complex structures are not always necessary
[8] G. E. Karniadakis, I. G. Kevrekidis, L. Lu, P. Perdikaris, S. Wang, and L. Yang, “Physics-informed machine learning,” Nat Rev Phys, vol. 3, no. 6, pp. 422–440, May 2021.
[9] C. W. Chang et al., “Dual-energy CT based mass density and relative stopping power estimation for proton therapy using physics-informed deep learning,” Physics in Medicine and Biology, vol. 67, no. 11, Jun. 2022
Conclusion
Efficiency of VMI in various E for both diagnosis and radiotherapy
VMI only reduces beam hardening artifact in metal artifact

## Slide 11
THANKS FOR YOUR ATTENTION

## Slide 12
APPENDIX
Evaluation image quality
Relative Difference:
x: pixels in testing images
y: pixels in target images

## Slide 13
APENDIX
Fig . Numerical simulation process.
Fig . Database construction for training.
DECT simulation and database preparation
Y. Zhang and H. Yu, “Convolutional Neural Network Based Metal Artifact Reduction in X-Ray Computed Tomography,” IEEE Transactions on Medical Imaging, vol. 37, no. 6, pp. 1370–1381, Jun. 2018