## Slide 1
Evaluation of HyperArc SRS Treatment Plan Quality Assurance Using Point Dose Measurements on Homogeneous and Heterogeneous Phantoms
STUDENT: TRƯƠNG GIA HỶ
HO CHI MINH City – 18/08/2025
IEEE NPSS SCHOOL ON ADVANCED TOPICS IN MEDICAL IMAGING

## Slide 2
1
Introduction
HyperArc SRS technique
QA of treatment plan
Device and measurement methods
OVERVIEW

## Slide 3
Research Objective
Perform point dose measurements at the irradiation isocenter using the HyperArc radiosurgery
Evaluate the accuracy and stability
Figure 1. Homogeneous phantom (SP34)
Figure 2. Heterogeneous phantom (STEEV)
2
OVERVIEW

## Slide 4
HyperArc SRS technique
3
HyperArc
SRS
Criteria
Intracranial treatment
Field arrangement
Automation
Dose distribution
Good, individual lesions
Optimized, for multiple lesions
Fixed, multiple directions
Non-coplanar arc arrangement
Manual treatment planning
Automated planning
Optimized, high conformity
Good
OVERVIEW

## Slide 5
SUBJECTS & METHODS
4
Subjects
31 patients with brain metastases treated with HyperArc.
Homogeneous phantom (Solid Phantom).
Heterogeneous phantom (Head-and-neck phantom).

## Slide 6
Figure 3. RAZOR Ionization Chamber
SUBJECTS & METHODS
Equipment
Figure 4. DOSE 1 Electrometer
Figure 5. TrueBeam 2.7 Linear Accelerator
5

## Slide 7
Figure 6. Homogeneous phantom with layered structure and detector attachment plate
Figure 7. Detector attachment plate compatible with the RAZOR ionization chamber
Homogeneous Phantom (Solid Phantom 34)
Figure 8. Heterogeneous phantom (STEEV)
Figure 9. Measurement module 038–03
Heterogeneous Phantom (Head-and-neck Phantom)
6

## Slide 8
Workflow of Treament Plan QA
Data collection
Treatment planning simulation on two phantoms
Data transfer from TPS to linear accelerator
Dose measurement and recording at the irradiation isocenter
Comparison, evaluation, and analysis of results
SUBJECTS & METHODS
7

## Slide 9
Figure 10. Patient Plan Data
Figure 11. Arcs and Fields in the Patient Plan
SUBJECTS & METHODS
Workflow of Treament Plan QA
8

## Slide 10
Figure 12. Treatment Plan Parameters
Figure 12.1. Dose Calculation on SP34
Figure 12.2. Dose Calculation on STEEV
Figure 12.1a. Dose Recording at the Reference Point (SP34)
Figure 12.2a. Dose Recording at the Reference Point (STEEV)
Workflow of Treament Plan QA
9

## Slide 11
Experimental Setup
10
Figure 13. Setup of SP34 phantom on the TrueBeam 2.7 linear accelerator
Figure 14. Setup of STEEV phantom on the TrueBeam 2.7 linear accelerator
SUBJECTS & METHODS

## Slide 12
Figure 15. SP34 Marked with the Laser Positioning System
Figure 16. STEEV Marked with the Laser Positioning System
SUBJECTS & METHODS
Experimental Setup
11

## Slide 13
Phantom Position Verification
Figure 17. CBCT Image of SP34
Figure 18. CBCT Image of STEEV
SUBJECTS & METHODS
12

## Slide 14
Results Evaluation
Dmeasured: Measured dose (Gy)
DTPS: Dose calculated by the treatment planning system (Gy)
): Dose deviation (%)
≤ 5% (AAPM TG-218)
SUBJECTS & METHODS
Acceptance Criteria
13

## Slide 15
EXPERIMENTAL RESULTS
No. | Patient | DPrescribed (Gy/frac) | DMeasured (Gy) | DTPS (Gy) | DD (%)
1 | NSH | 9 Gy/1 | 7,627 | 7,425 | 2,716
2 | LVT | 9 Gy/1 | 8,580 | 8,488 | 1,084
3 | PTTT | 9 Gy/1 | 6,697 | 6,619 | 1,173
4 | VTTH | 9 Gy/1 | 2,220 | 2,213 | 0,321
5 | NTK | 9 Gy/1 | 5,202 | 5,131 | 1,390
6 | NTLE | 6 Gy/1 | 2,425 | 2,376 | 2,070
7 | BHP | 24 Gy/1 | 21,926 | 21,005 | 4,383
8 | LTTT | 20 Gy/1 | 2,104 | 2,074 | 1,431
9 | VNH | 24 Gy/1 | 21,593 | 21,108 | 2,299
10 | BTHV | 6 Gy/1 | 5,390 | 5,299 | 1,711
11 | TVH | 20 Gy/1 | 17,648 | 17,305 | 1,980
12 | TTTN | 9 Gy/1 | 7,196 | 7,153 | 0,604
13 | NTD | 9 Gy/1 | 7,601 | 7,500 | 1,347
14 | TMT | 24 Gy/1 | 19,097 | 18,63 | 2,507
15 | LTL | 20 Gy/1 | 16,691 | 16,545 | 0,883
16 | LTKL | 9 Gy/1 | 7,699 | 7,653 | 0,605
No. | Patient | DPrescribed (Gy/frac) | DMeasured (Gy) | DTPS (Gy) | DD (%)
17 | DVD | 9 Gy/1 | 8,071 | 8,01 | 0,767
18 | NTTH | 6 Gy/1 | 2,636 | 2,599 | 1,427
19 | NTHG | 9 Gy/1 | 1,299 | 1,311 | -0,895
20 | QTD | 9 Gy/1 | 1,200 | 1,237 | -2,971
21 | VVV | 6 Gy/1 | 0,441 | 0,442 | -0,236
22 | TTH | 6 Gy/1 | 0,886 | 0,877 | 1,053
23 | NTVD | 9 Gy/1 | 8,038 | 8,008 | 0,377
24 | TVT | 6 Gy/1 | 5,636 | 5,637 | -0,023
25 | GTHO | 9 Gy/1 | 8,218 | 8,222 | -0,053
26 | NMC | 9 Gy/1 | 0,916 | 0,907 | 1,033
27 | TNL | 24 Gy/1 | 3,196 | 3,113 | 2,667
28 | LTKH | 9 Gy/1 | 8,095 | 7,968 | 1,599
29 | TMVQ | 9 Gy/1 | 8,440 | 8,42 | 0,243
30 | CVP | 6 Gy/1 | 1,035 | 1,026 | 0,859
31 | TMQ | 20 Gy/1 | 15,683 | 15,463 | 1,424
Table 1. Measured Dose Results on the Homogeneous Phantom
14

## Slide 16
15
Figure 19. Dose Correlation Graph between TPS and Ionization Chamber Measurement
Homogeneous Phantom
Remarks
31/31 cases met the 5% acceptance criterion
Mean deviation: 1,090%
Standard deviation: 1,285%
DTPS = 0,9740Dmeasured + 0,0756
EXPERIMENTAL RESULTS

## Slide 17
No. | Patient | DPrescribed (Gy/frac) | DMeasured (Gy) | DTPS (Gy) | DD (%)
1 | NSH | 9 Gy/1 | 10,124 | 9,869 | 2,585
2 | LVT | 9 Gy/1 | 10,821 | 10,831 | -0,089
3 | PTTT | 9 Gy/1 | 8,550 | 8,63 | -0,928
4 | VTTH | 9 Gy/1 | 2,858 | 2,871 | -0,459
5 | NTK | 9 Gy/1 | 6,168 | 6,084 | 1,386
6 | NTLE | 6 Gy/1 | 2,899 | 2,838 | 2,144
7 | BHP | 24 Gy/1 | 28,082 | 26,942 | 4,231
8 | LTTT | 20 Gy/1 | 2,777 | 2,782 | -0,176
9 | VNH | 24 Gy/1 | 27,825 | 27,399 | 1,554
10 | BTHV | 6 Gy/1 | 7,159 | 6,999 | 2,283
11 | TVH | 20 Gy/1 | 22,653 | 22,163 | 2,212
12 | TTTN | 9 Gy/1 | 9,507 | 9,434 | 0,770
13 | NTD | 9 Gy/1 | 9,810 | 9,678 | 1,363
14 | TMT | 24 Gy/1 | 24,145 | 23,501 | 2,741
15 | LTL | 20 Gy/1 | 21,489 | 21,197 | 1,377
16 | LTKL | 9 Gy/1 | 9,968 | 9,856 | 1,141
No. | Patient | DPrescribed (Gy/frac) | DMeasured(Gy) | DTPS (Gy) | DD(%)
17 | DVD | 9 Gy/1 | 10,389 | 10,241 | 1,448
18 | NTTH | 6 Gy/1 | 3,488 | 3,413 | 2,191
19 | NTHG | 9 Gy/1 | 1,578 | 1,584 | -0,359
20 | QTD | 9 Gy/1 | 1,443 | 1,498 | -3,657
21 | VVV | 6 Gy/1 | 0,531 | 0,532 | -0,123
22 | TTH | 6 Gy/1 | 1,145 | 1,122 | 2,075
23 | NTVD | 9 Gy/1 | 10,453 | 10,346 | 1,034
24 | TVT | 6 Gy/1 | 7,512 | 7,455 | 0,760
25 | GTHO | 9 Gy/1 | 11,038 | 10,936 | 0,936
26 | NMC | 9 Gy/1 | 1,127 | 1,142 | -1,334
27 | TNL | 24 Gy/1 | 4,220 | 4,019 | 5,000
28 | LTKH | 9 Gy/1 | 10,528 | 10,347 | 1,746
29 | TMVQ | 9 Gy/1 | 11,221 | 11,111 | 0,988
30 | CVP | 6 Gy/1 | 1,292 | 1,298 | -0,448
31 | TMQ | 20 Gy/1 | 20,191 | 19,883 | 1,551
16
Table 2. Measured Dose Results on the Heterogeneous Phantom
EXPERIMENTAL RESULTS

## Slide 18
Remarks
31/31 cases met the 5% acceptance criterion
Mean deviation: 1,095%
Standard deviation: 1,647%
DTPS = 0,9748Dmeasured + 0,0836
17
Figure 20. Dose Correlation Graph between TPS and Ionization Chamber Measurement
Heterogeneous Phantom
EXPERIMENTAL RESULTS

## Slide 19
18
Retrospective analysis of 31 brain metastasis cases
Comparison of dose deviations between TPS and actual measurements
  Homogeneous phantom: All 31/31 cases met the 5% criterion, with a mean deviation of 1.090% (0.023% – 4.383%).
  Heterogeneous phantom: All 31/31 cases met the 5% criterion, with a mean deviation of 1.095% (0.089% – 5.000%).
  The mean deviation between measured doses and TPS doses in the homogeneous phantom was significantly smaller compared to the heterogeneous phantom.
  A high linear correlation was observed between measured doses and TPS doses.
CONCLUSIONS