## Slide 1
Fang Liu
Institute of High Energy Physics, Chinese Academy of Sciences
A Bunch-by-Bunch Beam Position Monitor
中国科学院高能物理研究所
INSTITUTE OF HIGH ENERGY PHYSICS, CAS
2019-6-5

## Slide 2: Introduction
ADC Frequency &Accuracy Improved
Digital Signal Processing
Research&
Methodology
Deeping &Sufficient
High Precision
Bunch-by-Bunch Beam Position Monitor
Available
——Turn-by-Turn Measurement
——Fast Acquisition Measurement
——Slow Acquisition Measurement
doing Bunch-by-Bunch measurement
Based Research
Reserch of Inter-Bunch Beam Dynamics
Hardware platform of Bunch-by-Bunch Feedback system，Bunch-by-Bunch Beam Current Monitor；Diagnosing Beam Lost and Catching Tune of Beam
Normal BPM
Goal：Resolution of Bunch-by-Bunch : 5um

## Slide 3: A Bunch-by-Bunch BPM
4 Channels 500MSPS 14bit ADC
1GB DDR3
PCIE GEN2 X4
ZYNQ7035
100/1000Mbps ETHERNET
dual-core ARM Cortex-A9&Kintex®-7 ZYNQ7035
AD9684, 2 channels
Electronic System
A Four Button Pickup
Delay Module
The signals from the pickups of BEPCII are enough for directly data sampling of ADC
e+ , e-
Peak:300mV
Adjust to the time to guarantee ADC can get the peak of the Button Signal
Input Analog Band Width: 10MHz-2GHz
Notes: Using CIC and FIR to achieve non aliasing filtering and decimation.
Use the CORDIC module to calculate the magnitude
Use digital automatic gain controller to adjust the signal amplitude, so that let the FA and SA to get more effective bit

## Slide 4: Development state
2018
2019
Hardware：Circuit Design, Soldering and Debugging.
Firmware：ADC control, Clk control, DDR ,Ethernet
Bunch-by-Bunch Algorithm
2020
2021
Develop：Embedded Linux and EPICS on ZYNQ ARM
R&D Plan: A ZYNQ Based, Four channels 500Msps ADCs, Bunch-by-Bunch System
Complete
Already Down
Soft in ZYNQ ARM：Initialization of Ethernet, ADC and Clk. Data Acquisition Control.
Need to Do
Improvement: Upgrade Hardware and Firmware
Furthermore: Embedded Timing ?....

## Slide 5: Hardware
10/100/1000Mbps Ethernet
2 DDR3: one for Data Storage, one for ARM
PCIE GEN2 X4: backplane
ZYNQ7035 FPGA and 4 channels 500MSPS 14bit ADC
ZYNQ-7035
DDR3
PCIE
ETHERNET
CLK
TRIGGER
FMC
ADC1
ADC3
ADC2
ADC4
AD9684
AD9684
clk processing
CLK Management
Architecture
Base board and FMC Carrier
Notes: Using CIC and FIR to achieve non aliasing filtering and decimation.
Use the CORDIC module to calculate the magnitude
Use digital automatic gain controller to adjust the signal amplitude, so that let the FA and SA to get more effective bit

## Slide 6: Four Channel 500MSPS ADC FMC Card

## Slide 7: Base Board: ZYNQ7035

## Slide 8: ARM Core Construction

## Slide 9: FPGA Logical in ZYNQ PL

## Slide 10: Using PS code configure FPGA

## Slide 11: ARM Code in ZYNQ PS

## Slide 12: Test in Lab

## Slide 13: Test Environment
ADC CLK：500MHz （Convenient to synchronous two signal generator）
（CLK：From ROHDE&SCHWARZ SMA 100A 9KHz-3GHz）
Four Channel ADC Input：One to Four Power Splitter, Signal Source： (ROHDE&SCHWARZ SMF 100A 100KHz-22GHz)

## Slide 14: No Input, ADC Raw data
Peak to Peak:15
Maximum: ±8192 (14bit)

## Slide 15: 10MHz Input

## Slide 16: 490MHz Input

## Slide 17: 950MHz Input

## Slide 18: 1.5GHz Input

## Slide 19: 1.8GHz Input

## Slide 20: Summary
Develop：Embedded Linux and EPICS on ZYNQ ARM
Furthermore: Upgrading Hardware and firmware, adding embedded Timing module…
Improvement: Upgrade of Hardware and Firmware
Bunch-by-Bunch BPM system has been designed.
Prototype circuit board has been made.
Initial board tests have started.
Next Step:

## Slide 21
Thank you！