## Slide 1: Temperature Control of Crystal Optics The meV-resolved Inelastic X-ray spectrometer at NSLS-II
Kazimierz Gofron
EPICS Collaboration Meeting
Ljubljana, Slovenia
September 21, 2022

## Slide 2: Inelastic X-ray Spectrometer (9.1keV)
Back Row (L-R): Dima Bolmatov, Rick Greene, Scott Coburn, Kazimierz Gofron
Front Row (L-R): Alexey Suvorov, Alessandro Cunsolo, Ron Pindak, Misha Zhernenkov, Yong Cai
PRESENT AND FORMER TEAM MEMBERS
Montel Mirror
CDW
Sample
Analyzer Optics

## Slide 3: Optical Layout of Montel-CDW IXS Analyzer
Detector
Operation energy: 9.1 keV – Si(008) back reflection
Angular acceptance : ~ 7.5 (V) x15 (H) mrad2
200mm
~1.5 mm ~0.1 mrad
~10 µm
1.2 m (6 crystals of 200mm each)
~10 mrad Acceptance
Angular dispersive optics with post-sample collimation
(Montel)
Montel Mirror
~
Thickness ~ 200 um
Temperature < 1 mK
Projected Longitudinal Length @ 40 um Transerve Acceptance

## Slide 4: Inelastic X-ray beams from 6 D-crystals
Lambda (Medapix3) 2D detector imaging (right)
Temperature feedback of D-crystals (below)
Blade-type C crystal
D1
D3
D5
D2
D4
D6
With
W/O
D crystal φ de-tuning
Keithley 3706A
seq ioc
ePID

## Slide 5: Improving D Crystal Support with T Control
Supporting D crystals at their Airy points reduces lattice bending by gravity.
EPICS-based temperature control achieving stability : ± 1 mK/24hr.
Temperature of all D crystals adjusted to align and maintain diffraction energy and optimize reflectivity
D1
D2
D3
D4
D5
D6
Before
Energy aligned by RC adjustment
Airy Point Support
Attached heater
After
Temp Control
Current performance limited by crystal quality of the long D crystals.
Energy positions at peak of RC

## Slide 6: Temperature Response
Right: All 6 D-crystals control PT1000 sensors
Below: One D3-crystal fitted with 3 PT1000 sensors: control PT1000 sensor, and two PT1000 sensors at the crystal ends.

## Slide 7: Temperature stability
Required: < 4 mK / 1-week measurement
Before: 100-500 mK
After: ~0.1 mK
Scan: 60 sec
Control sphere | Stability | Contribution
Experimental Hall | +/- 1 K | Day/night cycle; Season and weather cycle
Hutch | +/- 0.5 K | Day/night cycle; Season and weather cycle
Chamber | +/- 0.2 K | Day/night cycle; Season and weather cycle
Crystal | +/- 0.0005K | N/A

## Slide 8: Summary
We have improved the data rate of the meV-IXS spectrometer by more than 13 times and a routine energy resolution at 1.7 meV with sharp tails since 2019, allowing soft matter research to be performed with much enhanced count rate and contrast and enabling science previously impossible (notably, on hard condensed matter).
  Temperature stability of the crystal optics analyzer played important role in stability.
Future: Stabilize temperature of crystals for High Resolution Monochromator, using identical system components.