## Slide 1
The prototype of high stiffness load cell for Rockwell hardness testing machine calibration according to ISO 6508-2:2015
“I AM NIMT”
M Pakkratoke1 and T Sanponpute2
1, 2 National Institute of Metrology (Thailand), 3/4-5 Moo 3, Klong 5, Klong Luang, Pathumthani, Thailand 12120 email: 1montree@nimt.or.th, 2tassanai@nimt.or.th

## Slide 2: Outline
Introduction to the verification of hardness machine
Force verification of hardness testing machine
Conceptual design of High Stiffness Load Cell
Experiment results
Conclusion
“I AM NIMT”

## Slide 3: 1. Introduction to direct verification of Rockwell hardness testing machine
2. Verification of the testing cycle
1. Verification of the test force
3. Verification of the depth measuring device
4. Verification of the indenter shape
Test Force
Time
Depth
Indenter shape
Hardness Machine
“I AM NIMT”

## Slide 4: Operation of hardness testing machine
Weight
Bearing
Rockwell hardness
Calibration position
“I AM NIMT”
Calibration position

## Slide 5: 2. Force verification
“I AM NIMT”
Each test force shall be measured.
Shall be done at not less than three position of the plunger movement during test.
Three readings shall be taken for each force at each position of the plunger.
The preliminary test force, F0 shall be measured before application and after removal of the additional test force, F1.

## Slide 6
Test force of Rockwell hardness testing machine
Rockwell hardness scale | Preliminary Test force, F0 (kgf) | Total test force, F (kgf)
15N, 15T | 3 | 15
30N, 30T |  | 30
45N, 45T |  | 45
A, F, H | 10 | 60
B, D, E |  | 100
C, G, K |  | 150
Preliminary Test force, F0
Total test force, F
Calibration position of Rockwell A | 
Hardness | Position
HRA | (µm/HR )
85 | -30/-15
60 | -80/-40
30 | -140/-70
Notes: Depth =-2*(100-HR)

## Slide 7
Calibration position changes due to load cell deformation
Weight
Bearing
Load cell deforms
Calibration position
The calibration position must be compensated!
“I AM NIMT”
Reference position
(HR)
100
Position after applied F1
(HR)
55
Deforms of load cell | 
( HR ) | ( µm )
100-55 = 45 | 452=90
Commercial load cell usually deforms more than 200 µm at capacity

## Slide 8
Capacity of load cell over 150 kgf
Deform less than 20 m
Compact size, suit for hardness machine
“I AM NIMT”
3. Conceptual design

## Slide 9: Design of HSL
“I AM NIMT”
A
B
C
D
4 sensing body =300 kgf

## Slide 10
“I AM NIMT”
Parameters | PA | PB | PC | PD
Stress (GPa) | 117.5 | 127.4 | 116.7 | 127.9
Strain (×10-6 strain) | 458.0 | 501.3 | 457.9 | 498.7
Yield strength = 350 GPa
PA
PB
PD
PC
FEA simulation results
150 kgf applied force
4. Experiment results

## Slide 11
“I AM NIMT”
Parameters | A | B | C | D | E | F | G
Deformed (µm) | 0.95 | 6.56 | 6.68 | 6.16 | 8.0 | 8.0 | 9.25
Calculate (mm) | - | - | - | - | - | - | 5.00
Displacement simulation results
A
B
C
D
E
F
G
Weakness of pressure plate

## Slide 12
Fabricated HSL
“I AM NIMT”
Dimension size of a fabricated is slightly different from design.

## Slide 13
“I AM NIMT”
Deformation experiment results
Test force (kgf) | HSL experiment results (µm) | Simulation results (µm)
15 | 2.4 | 0.9
30 | 4.4 | 1.8
45 | 5.9 | 2.8
60 | 5.6 | 3.7
100 | 8.5 | 6.2
150 | 12.1 | 9.2
The deformation experiment results are only 2 µm larger than FEA results.

## Slide 14: Conclusion
The 150 kgf high stiffness load cell was designed and constructed.
The deformation under 150 kgf is lower than 10 µm.
The strain under 150 kgf is 542.57×10-6 m/m.
FEA and experiments gave the consistent results within 5 µm displacement length and 50×10-6 m/m strain.
Weakness of pressure plate.
Load cell calibration and classification.
Future work