## Slide 1: Online Optical Probes for Quality Control and Safety Assessment of Olive and Other Edible Oils
G. Stavropoulos
Demokritos, November 2013

## Slide 2
Sample
Spectrophotometer
Grating
Light Source
Light Sensor
Collimating Optics
Optical fibres

## Slide 3
Incoming light optical fibre
Adjustable slit
Diffraction grating
Readout electronics
SPECTROPHOTOMETER PROTOTYPE
Linear CMOS sensor (1024 pixels)

## Slide 4
SENSOR
MODULE
SIGNAL PROCESSING
A/D CONVERTER
SPI BUS
MICROPROCESOR UNIT
USB, Wi-Fi BLUETOOTH BUS
USER INTERFACE
CONTROL BUS

## Slide 5: IN HOUSE CMOS MODULE
After designing the CMOS board a set of tests took place for the linearity and dark RMS noise of the sensor-board module.
We used a modified wooden box to place the module and another case for the logarithmic filter and the single axis translation stage.
Modified cases where developed by G.Kiskiras
and D.Vakondios
Measurements where analyzed with a Linux software developed by D.Lenis

## Slide 6: TEST CONDITIONS
CMOS BOARD
OPTIC FIBER

## Slide 7: TEST CONDITIONS
Logarithmic filter
TURN SCREW

## Slide 10: CMOS PHOTO OF INCOMING LIGHT FROM FIBER

## Slide 11: DARK RMS
29 counts=1,5 mV

## Slide 12
2nd order diffraction of 253nm line
Light source output
CMOS output
Wavelength (nm)►
Example measurement set (32 measurements) using calibration source

## Slide 13
Detail picture of the 253nm peak for the CMOS readout

## Slide 14: PROBEOIL software Σωτηρία Δήμου
Απαιτήσεις λογισμικού για την συσκευή
Δημιουργία interface σε γλώσσα προγραμματισμού JAVA
Σύνδεση με την συσκευή
Λήψη δεδομένων από την συσκευή (128, 512 η 1024 αριθμούς με endline στο τέλος του κάθε αριθμού)
Αποθήκευση δεδομένων σε αρχείο
Αλγόριθμος - πολλαπλασιασμός με μια σταθερά για τον υπολογισμό τιμών έντασης φωτός
Αποθήκευση νέων δεδομένων σε αρχείο
Μελλοντικά – Δημιουργία Διαγράμματος τιμών έντασης φωτός
ανά μήκος κύματος

## Slide 15: Το interface
Connect – Σύνδεση με την θύρα που έχουμε επιλέξει και λήψη δεδομένων μέσω JAVA και αποθήκευση τους σε ένα αρχείο
Get data from port – Σύνδεση με την θύρα που έχουμε επιλέξει και λήψη δεδομένων μέσω C και αποθήκευση τους σε ένα αρχείο
Calculate Data – Αλγόριθμος για τον υπολογισμό τιμών έντασης φωτός και αποθήκευση τους σε ένα αρχείο
Disconnect – Αποσύνδεση από την θύρα και αποδέσμευση της

## Slide 16: CCD SENSOR MODULE
CCD 3648 pixels
Spectral response 200-900 nm
Low noise PCB 6-Layer Design
FPGA based Data Acquisition
Under development by T.Athanasopoulos

## Slide 17: APD MODULE

## Slide 18: IN HOUSE APD MODULE
High Voltage module
Temperature control module
Signal processing(current to voltage converter,
A/D converter)
Step motor module
Motherboard to control all peripherals

## Slide 19: HIGH VOLTAGE MODULE
Resonant Royer Based DC/DC Converter with controlled feedback
Controlled Output Voltage 0-250 V
Very low ripple < 0,01 % V p-p
Low peak on rise time
Under development by M.Maniatis, T.Domvoglou

## Slide 20: TEMPERATURE CONTROL MODULE
Thermo electric cooler (TEC)
Current control circuit for TEC
Thermo Sensor TO-220 casing
8-bit Slave Microcontroller
Under development by T.Domvoglou

## Slide 21: SIGNAL PROCESSING
Current to Voltage circuit
16-bit A/D Converter
Low noise PCB
Developed by M.Maniatis, T.Domvoglou

## Slide 22: STEP MOTOR MODULE
Low turn step motor < 2 degrees
Step on X-Axis 0,5-3 mm
Accurate positioning
Under development by M.Maniatis

## Slide 23: Next steps
Start sample measurements with the first prototype to evaluate its performance
Improve the mechanics of the prototype
Evaluate gratings with more grooves/mm for better wavelength accuracy (ordered)
Create the first prototype for NIR spectroscopy (mechanical electronical and optical components have been ordered)
  InGaAs photodiode 900-2500nm response
  Pds photodiode 1000-2900nm response
Evaluation and development of alternative sensor technologies
  APD based module (under development)
  CCD based module (under development)
  Use a CMOS sensor with more pixels (under development)
  PMT module 200-800nm response
  MPPC module 320-900nm response