## Slide 1: Developments in 2D XRF analysis for Macro scale imaging of Cultural Heritage artefacts
Vasiliki Kantarelou, Claudia Caliri, Andreas Germanos Karydas and
Francesco Paolo Romano
Institute of Nuclear and Particle Physics, N.C.S.R. “Demokritos”
Laboratori Nazionali del Sud, INFN

## Slide 2: Schematic representation of the various modes in XRF analysis
Milli XRF
Beam size ~ mm
Old Micro XRF
Beam size : 20-100 μm
2D
Max analyzed area
~10 x 10 mm2
Confocal XRF
Volume ~
20x20x20 μm3
3D
Macro area Micro XRF
Beam size : 20 μm-1 mm
2D
Analysis of 110 x 70 cm2 in 4.3 h
with continuous scanning 100 mm/s

## Slide 3: Micro-XRF spectrometer
X-ray Rh-anode tube (spot size 50×50 μm, max 50 kV, max 1 mA, 30 W maximum power consumption, Be window 0.2-mm thickness)
Polycapillary X-ray lens (IFG), focal distance 21.2 mm.
Silicon drift X-ray detector (10-mm2) with 146 eV FWHM at 10 kcps coupled with a digital signal processor.
Color CCD camera (×13), dimmable white LED, laser
Three stepping motors
Customized filter composed by
Ti (23.6 ± 0.2) μm and Co (17.7 ± 1.3) μm
Spatial resolution: 40 - 80 μm for characteristic X-rays between 22-8 keV, respectively.
A customized version of ArtTAX, (Bruker Nano GmbH)

## Slide 4: MA-XRF spectrometer
X-ray source: Rh target (by IFG)
Polycapillary optic with a focus size of 25 μm at 1cm (by IFG)
Two 50 mm2 SDDs with 130 eV @ 5.9 keV (by KETEK)
Digital X-ray processors operated in time list event mode with 40 ns time resolution (by BrightSpec)

## Slide 5
The continuous stream of X-ray events is elaborated by the
Control Unit as coming from a single detector
SDD 2
SDD 1
Scanning with a uniform continuous speed up to 100 mm/sec
T11 E11
T12 E12
.
.
T1n E1n
T21 E21
T22 E22
.
.
T2n E2n
T11 E11
T21 E21
T12 E12
T22 E22
.
T1n E1n
T2n E2n
X-ray spectra are obtained by packing data in pixels of dimension fixed by users. Pixel size can be changed dynamically during scanning
5300 fitted spectra per second
A Control Unit (CU) programmed under Real Time Linux is used to control all sensors in the scanner, to acquire data and elaborate X-ray spectra in a live mode
Real-time fitting of pixel XRF spectra

## Slide 6
Scanning speed @ 100 mm/sec
Elemental maps elaborated with 1 mm pixel size (equivalent to 10 ms dwell time)
Pixels: 3.2 Mega; Scanning time 9.5 h (including a full fitting of pixel spectra)
Mocking of Christ, 1640 ca.,
Matthias Stomer (210x154 cm2)
Fast MA-XRF scanning

## Slide 7: Micro-XRF quantification methodology
Determination of transmission efficiency using a set of thin and thick targets
Validation of the method using certified reference gold, bronze and silver alloy and glasses using the ratio of measured to certified concentration (Ki)
Glasses
Metals
V. Kantarelou and A.G. Karydas, XRS, 2016, 45, 85-91.

## Slide 8: Micro-XRF application: Gold signet rings LBA
SN24-18
SN24-30
SN24-702
SN24-736
The rings were recovered in 2015 in the grave of the Griffin Warrior at the Palace of Nestor in Pylos by the archaeologists of the Cincinnati university.

## Slide 9: Micro-XRF application: Gold signet rings LBA
Cu content is below 4 %
Use of native gold
Range of Au : 75 - 92.5 %
Range of Ag : 7.5 – 27.5 %

## Slide 10: MA-XRF application: Pitsa wooden panels, Pinax A
MA-XRF mapping
White tone:
High intensity of an element
Grey tones:
Low intensity of an element
The white preparatory substrate is identified as gypsum (CaSO4) that must have been mixed with an organic binder, presumably animal glue, in order to give strength to the ground and fix it to the wooden surface

## Slide 11: MA-XRF application : Pitsa wooden panels, Pinax A
16 November 2017
11
Dark red of the chiton and himatia: Hematite (Fe2O3)
Men's’ skin:
Mixture of hematite and gypsum (CaSO4)
Blue of the himatia: Egyptian blue
CaCuSi4O10
Fe: Red
As: Green
Cu: Blue

## Slide 12
Thickness ~ 10 μm
Ochre: 53.6%
Binder: 45%, Trace elements: Cu, As, Sr
MA XRF Quantification: Pinax A
Ochre Layer
Gypsum: 75.2%, Binder: 21.4%
Lime: 1.8%, K2O: 1.0%
Trace elements: Ti, Fe, Cu
Gypsum
Notes: We performed a Monte Carlo simulation on the red ochre layer so we calculated the thickness of the layer as well as the concentration of the pigment (ochre) and the binder.

## Slide 13: Future plans
Build our own Macro-XRF
Quantitative analysis of MA-XRF using Monte Carlo XMI-MSIM based simulations
Thank you for your attention !