## Slide 1: A modular custom-built Macroscopic-XRF (MA-XRF) spectrometer for Heritage Science
Kalliopi Tsampa (1), Effrossyni Androulakaki (1), Panagiotis Assiouras (1), Paweł Wróbel (2)
and Andreas G. Karydas (1)(3)
Institute of Nuclear and Particle Physics, NCSR “Demokritos”, 153 41, Agia Paraskevi, Attiki, Greece
(2) Faculty of Physics and Applied Computer Science, AGH University of Science and Technology, al. Mickiewicza 30, 30-059 Kraków, Poland
(3) CNR, Istituto per i Beni Archeologici e Monumentali (IBAM), Via Biblioteca 4, 95124, Catania, Italy
National Center for Scientific Research “Demokritos”
Institute of Nuclear and Particle Physics

## Slide 2: Aim of the research
Analytical characterization of the excitation features of the MA-XRF spectrometer by means of Monte Carlo simulations tools and dedicated experimental measurements.
More specifically,
Simulation of end window tube spectrum - PenRed code
Simulation of the optical features of X-ray polycapillary lens – Polycapillary X-ray raytracing Polycap
Under different setup configurations, experimental evaluation of the spectrometer
  Intensity through output
  Spatial resolution

## Slide 3: The spectrometer instrumentation
X-ray beam

## Slide 4: The spectrometer instrumentation
X-ray beam
X-ray Detector
Laser triangulation sensor for the sample distance measurement
X-ray tube
Detector shielding
X-ray tube shielding
Arduino (homebuilt) for the shutter control
X-ray Detector SDD 150mm2
Digital Signal Processor DANTE
X-ray Tube
12W (60kV, 200uA) MagPro end window Rh-target X-ray source, 400um spot size
Laser sensor
Motorized XYZ Stage
750mm x 750mm x 250mm

## Slide 5: The spectrometer modularity
Polycapillary lens
Pinhole configuration
Fixed position of the SDD
Adjustable position of the X-ray tube
Identical focus position for pinhole and lens
Polycapillary Lens configuration
Pinholes
0.5mm, 1mm, 4mm
Polycapillary X-ray Lens
FWHM @10 keV ~ 87um
Pinholes

## Slide 6: Custom-made LabView acquisition controller
  Features:
  Acquisition parameters and hardware control: velocity, acceleration of the stages, X-ray tube shutter, proximity sensor
  Stable velocity during acquisition
  minimize possible shifts
  typical scanning 6 hours for 50 x 50 cm2
  Spot and scanning measurements with
  stepping or continuous “free-running” mode
  Elemental maps (up to 4 elements) using ROIs during acquisition

## Slide 7: Excitation spectrum – Monte Carlo simulation
* (Giménez-Alventosa et al., 2021)
MC simulation of end window Rh anode spectrum
Monte Carlo code PenRed *
A general-purpose, stand-alone code based on PENELOPE for Monte Carlo simulations of electron-photon transport, allows for parallel processing drastically reducing the computing time
Simulation parameters:
Tube high voltage
Thickness of the Rh anode
Indicatively, MC results are compared to Ebel semi-empirical model
Computing time: 1-3 days depending on input file using multi-threading capabilities (parallel simulation using 15-40 threads), statistics ~7% - ~ 25% for low/high photon energies

## Slide 8
X-ray tube excitation spectrum using Monte Carlo (MC) simulation
case 1: side window Rh tube – 40 kV
Comparison of MC simulation (PenRed) with the Ebel semi-empirical model yielded satisfactory agreement
MC results are compared to semi-empirical model.
The Ebel semi-empirical model is validated for side window (thick anode material) tubes. Regarding end window tubes, this model is not validated.
Comparison for a side – window Rh target tube

## Slide 9
Comparison of continuous spectrum for a transmission tube 0.75um Rh target
Semi-empirical model not validated for end window (transmission) X-ray tube
15kV
40kV
X-ray tube excitation spectrum using Monte Carlo (MC) simulation
case 2: end window Rh tube

## Slide 10
Comparison of discrete spectrum for 0.75um Rh target tube – 15 to 50 kV
Rh – Kα lines
Rh – L lines
The semi-empirical model overestimates significantly the discrete intensity at high voltages
Simulation evaluation with corresponding experimental data
03/29/2024
Tsampa Kalliopi – INPP Annual Meeting

## Slide 11: X-ray tube excitation spectrum - Experimental
Length: 1m
Collimators: Cu/Sn
Voltages: 12 - 55kV
Comparison to PenRed code (in progress)
known solid angle
X-ray tube
X-ray Detector
Collimators
Vacuum
top
top view

## Slide 12: Monte Carlo based ray-tracing simulation –POLYCAP code*
* (Tack et al., 2020)
Monte Carlo based polycapillary X-ray tracing code: POLYCAP1.1*
simulation of photon propagation through lens
features surface roughness
includes capillary wall transmission (important in high energies)
includes beam polarization effects
POLYCAP1.1 code produces data at a certain energy for
FWHM of the beam

## Slide 13
f1: 55mm
f2: 24mm
Beam diameter at entrance
(6 mm)
Beam diameter at focus
(< 100 μm)
Polycapillary Lens

## Slide 14: POLYCAP vs Experiment: Beam size @ focus
Comparison to nominal values
Tube spot size (400μm)
Microbeam tube (50μm)
Comparison to experimental data

## Slide 15
POLYCAP vs Experiment: Beam size out of focus
exp FWHM @ focus (μm) | 
Ti-Kα | 120 ± 2
Cu-Kα | 85 ± 1
Μο-Κα | 41 ± 2
Comparison to experimental data
Experimental data

## Slide 16
Experiment: Beam size at / out of focus
Cu grid – mesh 50
Hole width (μm): 420
Pitch (μm): 500
Bar width (μm): 80
Rim width (mm): 0.225
Thickness (μm): 20±3
at focus
3 mm away from focus
Dwell time/pixel: 100ms
Step: 50μm

## Slide 17: First applications
The main objective of the project is the development of an Open-Access Workshop (OAW) within the premises of the National Gallery at Athens (NGA), where analytical systems based in imaging and spectroscopic techniques will be installed for study of paintings.

## Slide 18
Painting “30th March 1814” – Scanned Areas

## Slide 19: Painting “30th March 1814” – Elemental maps
Ca
Fe
Bone Black (?)
Red ochre, Mars Black (?), ….
Co
Hg
Cobalt Blue
Cinnabar
Scanning of an area:
26.7 x 13 cm2
total scanning time: ~1h

## Slide 20: Painting “30th March 1814” – Composite elemental maps
Ca
Co
Hg
Fe
The distribution of elements can be visualized with different pseudo-colors
Their combination gives information on the composite distribution of the elements
Ca
Fe
Co
Hg
Scanning of an area:
26.7 x 13 cm2
total scanning time: ~1h

## Slide 21: Painting “30th March 1814” – Elemental maps
Ca
Cr
Fe
Red ochre,
Mars Black (?), ….
Viridian (?)
Bone Black (?)
Scanning of an area:
17 x 16.5 cm2
total scanning time: ~2h
Co
Hg
Cobalt Blue
Cinnabar

## Slide 22: Painting “30th March 1814” – Composite Elemental maps
The distribution of elements can be visualized with different pseudo-colors
Their combination gives information on the composite distribution of the elements
Ca
Co
Cr
Fe
Hg
Ca
Cr
Fe
Co
Hg
Scanning of an area:
17 x 16.5 cm2
total scanning time: ~2h

## Slide 23
Thank you for your attention!
National Center for Scientific Research “Demokritos”
Institute of Nuclear and Particle Physics